Supporting the Moldovan authorities in the sustainable management of the Dniester River

Research Monitoring Programme for Identification of WFD  ‘Priority and certain other pollutants’ in the Moldovan Dniester River Basin District

Final Report

Table of Contents

1	Introduction	1

2	Approach	3

2.1	Sampling	3

2.1.1	Sampling sites	3

2.1.2	Sampling scheme	4

2.1.3	Sampling	4

2.2	Laboratory analysis	5

2.2.1	Analysed substances	5

2.2.2	Laboratory equipment and methods of analysis	6

2.2.3	Quality Assurance & Quality Control (QA/QC)	8

3	Results	10

3.1	In situ measured parameters	10

3.2	Priority substances not detected	10

3.3	Detected Priority substances	11

3.3.1	AA-EQS, MAC-EQS	15

3.3.2	Heavy metals	15

3.3.3	Organochlorine pesticides	16

3.3.4	Other pesticides	16

3.3.5	Polyaromatic hydrocarbons	17

3.3.6	Volatile organic compounds	17

3.3.7	Industrial pollutants	18

3.3.8	Tributyltin compounds	18

3.3.9	Perfluorinated compounds, PFOS	18

3.3.10	Per- and poly-fluorinated alkyl substances, PFAS	19

4	Discussion	20

4.1	Priority substances not detected	20

4.1.1	Sample media	20

4.1.2	LOD > AA-EQS	20

4.1.3	Extrapolation	20

4.2	Priority substances detected	21

4.2.1	LOD, LOQ	21

4.2.2	Substances detected at ‘critical levels’: MAC-EQS	22

4.2.3	Substances detected at ‘near critical levels’: AA-EQS	22

4.2.4	Not detected at all sites	23

4.2.5	Not at all times	23

4.2.6	Hydrological considerations: Bălţi – Ustia	24

4.3	Synthesis	24

4.3.1	Capacity development of the Water Quality Laboratory	25

4.3.2	Design of surface water quality monitoring programmes	26

5	Conclusions	28

Abbreviations and acronyms

| AA-EQS | annual average environmental quality standard |
| --- | --- |
| EQS | environmental quality standard |
| GC | gas chromatograph |
| GC-MS/MS | gas chromatography tandem mass spectrometry |
| GC-MS(NCI) | gas chromatography mass spectrometry, negative ion chemical ionization |
| HM | heavy metals |
| ICP-MS | inductively coupled plasma mass spectrometry |
| IND | industrial pollutants |
| LC | liquid chromatography |
| LC-ESI-MS/MS | liquid chromatography-electrospray ionization-mass spectrometry |
| LC-ESI-QTOF-MS | high-performance liquid chromatography coupled with electrospray ionization quadrupole time-of-flight mass spectrometry |
| LC-MS/MS | liquid chromatography tandem mass spectrometry |
| LC-HRMS | liquid chromatography high resolution mass spectrometry |
| LOD | limit of detection |
| LOQ | limit of quantification |
| MAC-EQS | maximum allowable concentration environmental quality standard |
| MOLDAC | National Accreditation Centre of the Republic of Moldova |
| OCP | organochlorine pesticide |
| PAH | polycyclic hydrocarbons |
| PEST | other pesticides |
| PFAS | per- and poly-fluorinated alkyl substances |
| PFOS | perfluorinated compounds |
| QA/QC | quality assurance & quality control |
| RBMD | river basin management district |
| TBT | tributyltin compounds |
| VOC | volatile organic compound |
| WFD | Water Framework Directive |

# Introduction

As a candidate EU Member State, the Republic of Moldova is getting itself prepared for complying with the acquis communautaire. The Moldova-European Union Association Agreement (signed on 27 June 2014) came fully into force on 1 July 2016.

The UNDP project “Supporting the Moldovan authorities in the sustainable management of the Dniester River” provides – inter alia – support towards reaching compliance with the Water Framework Directive, including legal harmonization, strengthened water quality monitoring, and improved data processing and reporting in line with EU integration obligations.

Chapter 27 of the Association Agreement deals with environment and climate change, among which the Water Framework Directive (WFD) . The WFD concerns a comprehensive approach towards integrated management of the quality of surface waters and groundwaters, with ‘good status of all bodies of waters’ as ultimate objective.

For surface water this objective means that the status of water bodies must of good ecological status and good chemical status.

- The ecological status is largely determined by biological quality elements, such as fish, water plants, organisms dwelling in and on the river bed, phytoplankton, and more.

- The chemical status is determined by the so-called ‘Priority and certain other pollutants’, currently defined in the Directive 2013/39/EU .

The Project Document refers to chemical status under

- COMPONENT 4: Implementation of the on-the-ground activities to improve current ecological conditions of the Dniester River

  - Outcome 8: Chemical status of water bodies in the Dniester River Basin District

    - Output 8: Sampling program on Dniester, tributaries, and underground waters implemented, and data obtained and analysed.

The Project Document notices that little information is available about the presence of the ‘Priority and certain other pollutants’ in Moldovan surface waters. Therefore, it instigated to conduct sampling and laboratory analysis under a ‘Research monitoring programme’. This programme was designed to comprise 105 sites, to be visited once per year during three years, with water, sediment and biota samples to be taken. The samples were to be analysed the Water Quality Laboratory of the Environmental Reference Laboratory of the Environment Agency.

However, during a scoping mission in September 2024 it turned out that the Water Quality Laboratory was not capable of analysing the full range of ‘Priority substances and certain other pollutants’ (hereinafter referred to as Priority substances).

Preliminary orientations among laboratories abroad indicated that costs for analysis for Priority substances in one water sample amounted to around USD 2,000. Implying costs for analysis of 105 (sites) * 3 (years) water samples of about USD 600,000 (excluding various overhead); disproportionally high compared with the then project budget of USD 2.1 million.

Thus, the research monitoring programme had to be revised. In advance it should be emphasized that the main purpose of the research monitoring programme is to examine the presence/absence of Priority substances. Both the number of sites (9), as well as the number of samples taken per site (3) are in principle too small for assessment of the chemical status of delineated surface waterbodies inside the Moldovan Dniester River Basin District.

The programme was approved by the Project Board during their meeting in October 2024.

The related Request for Quotation was published in December 2024. A suitable offer was submitted by the Environmental Institute from Koš, Slovak Republic, taking care of the sampling and the laboratory analysis.

The Final Report of the Environmental Institute is issued separately. Basic information about sampling and laboratory analysis has been inserted into Chapter 2 but that report also addresses, inter alia, obervations and recommendations that are relevant for the Water Quality Laboratory.

# Approach

## Sampling

### Sampling sites

Nine sites were selected, presumed to be representative for wider areas of the Moldovan Dniester River Basin Management District (RBMD). Most sites are included in the regular national surface water quality monitoring programme, except for Nimereuca and Mereneşti.

Table 2-1 Features of the initially scheduled sampling sites

| River | Location | Features |
| --- | --- | --- |
| Dniester | Otaci | Upstream border with Ukraine |
|  | Nimereuca | End of stretch with shared MD-UA border; entry into Moldovan territory |
|  | Palanca | Downstream border with Ukraine |
| Răut | Bălţi | Downstream large city |
|  | Ustia | Downstream overall: combination of point sources and agriculture |
| Ichel | Goian | Downstream overall: combination of point sources and agriculture |
| Bîc | Străşeni | Downstream medium-sized city; agriculture (and upstream Ghidighici Reservoir; kind of watershed) |
|  | Gura Bîcului | Downstream large city |
| Botna | Mereneşti | Downstream overall: combination of point sources and agriculture |

There were deviations from the scheduled programme at two sites during the implementation:

- Dniester, Otaci: due to external factors, the April 2025 samples were taken a few hundred meters further downstream at the right bank, instead of from the bridge.

- Botna, Mereneşti: for practical reasons it was decided to move the sampling site along the Botna River to Chircăiești, 6 kilometres upstream.

The sampling sites are plotted in Figure 2-1 further below and their GPS coordinates in Table 2-2.

Table 2-2 Coordinates of the actual sampling sites

| River | Location | Latitude | Longitude |
| --- | --- | --- | --- |
| Dniester | Otaci (April) | 48.434278 | 27.804278 |
|  | Otaci (July, October) | 48.443993 | 27.790373 |
|  | Nimereuca | 48.116734 | 28.542365 |
|  | Palanca | 46.421673 | 30.128176 |
| Răut | Bălţi | 47.758995 | 27.961387 |
|  | Ustia | 47.252619 | 29.133935 |
| Ichel | Goian | 47.131857 | 28.912166 |
| Bîc | Străşeni | 47.136695 | 28.663687 |
|  | Gura Bîcului | 46.913183 | 29.457824 |
| Botna | Chircăiești | 46.719348 | 29.557359 |

Figure 2-1 Overview of the sampling sites

### Sampling scheme

The samples were taken during 18-20 April 2025, 14-16 July 2025 and 16-18 October 2025, following the generic schedule in Table 2-3.

Table 2-3 Sampling scheme per survey round

| Day | Site |
| --- | --- |
| Day 1 | Otaci |
|  | Nimereuca |
|  | Bălți |
| Day 2 | Ustia |
|  | Goian |
|  | Străşeni |
|  | Gura Bîcului |
|  | Chircăiești |
| Day 3 | Palanca |

### Sampling

The water samples were taken with buckets or, where feasible, directly in the bottles via submerging them at the water surface.

- Two 2 l Nalgene bottles , one 100 ml glass VOC bottle and one 100 ml PET bottle for metals were collected from each sampling point. The Nalgene and VOC bottles were prepared in the laboratory before each campaign - rinsed with millipore water, then acid cleaned (soaked in 2% nitric acid for 48 h), then rinsed with ultrapure (LC-MS grade) water, and then with LC-MS grade methanol. PET bottles were rinsed with millipore water, then rinsed with ultrapure (LC-MS grade) water and then with methanol (LC-MS grade).

- The water sample for determination of metals was obtained by filtration through a 0.45 µm filter, directly at the sampling point. The EQS for metals refers to the dissolved concentration (100 ml water filtered through syringe with Luer connector).

- A multiparameter portable device was used for measuring pH, temperature, conductivity, and oxygen content.

- The total hardness of the water was determined using a photometric kit.

## Laboratory analysis

### Analysed substances

Annex 1 contains the list with Priority substances and their environmental quality standards (EQSs) as included in the Directive 2013/39/EU. It concerns 45 main entries, some of them encompassing more than one substance.

Entry (37) Dioxins and dioxin-like compounds was omitted from laboratory analysis since no EQSs for water are defined in the Directive 2013/39/EU. Furthermore, its analysis is rather expensive.

The Environmental Institute also analysed and reported PFAS derivatives. PFAS as such are not included in the Directive 2013/39/EU, although PFOS are a subgroup. However, the following proposal is under preparation that includes PFAS:

Proposal for a directive of the European Parliament and of the Council amending Directive 2000/60/EC establishing a framework for Community action in the field of water policy, Directive 2006/118/EC on the protection of groundwater against pollution and deterioration and Directive 2008/105/EC on environmental quality standards in the field of water policy.

Table 2-4 provides an overview of the various categories of analysed Priority substances.

Table 2-4 Overview of categories of analysed substances

| Category | Acronym | Substances | Remarks |
| --- | --- | --- | --- |
| Heavy metals | HM | cadmium, lead, mercury, nickel |  |
| Organo chlorine pesticides | OCP | aldrin, chlorfenvinphos, DDT, more |  |
| Other pesticides | PEST | atrazin, diuron, quinoxyfen, more |  |
| Polyaromatic hydrocarbons | PAH | anthracene, fluoranthene, more |  |
| Volatile organic compounds | VOC | 1,2-dichloroethane, benzene, more |  |
| Industrial pollutants | IND | chloroalkanes, nonylphenols, more |  |
| Tributyltin compounds | TBT | tributyltin compounds |  |
| PFOS | PFOS | perfluorinated compounds | PFOS are a subgroup of PFAS |
| PFAS | PFAS | per- and polyfluoroalkyl substances | Not included in Directive 2013/39/EU |

The Environmental Institute reported a few more substances than explicitly mentioned in the Directive 2013/39/EU, which though can be considered as ‘family members’ of the Priority substances:

- DDD and DDE, which are part of total DDT.

- Aniline, styrene and toluene, being volatile organic compounds (VOCs) but not listed among the Priority substances.

- Linear perfluorooctanesulfonic acid, which is part of Total perfluorooctanesulfonic acid (PFOS).

In total 107 substances were analysed, among which 38 PFAS derivatives.

PFAS will be addressed separately in ensuing chapters and annexes, since they are not defined in the Directive 2013/39/EU.

### Laboratory equipment and methods of analysis

The methods of analysis are enumerated in Table 2-5.

Table 2-5 Methods of analysis

| Name of group | Substances in group | Method of analysis |
| --- | --- | --- |
| LC- Pesticides | aclonifen, alachlor, atrazine, bifenox, cybutryne, cypermethrin, diuron, Isoproturon, quinoxyfen, simazine, terbutryn | In-house modified STN EN ISO 6468 |
| GC-PAHs | anthracene, benzo(a)pyrene, benzo(b)fluoranthene, benzo(g,h,i)perylene, benzo(k)fluoranthene, fluoranthene, indeno(1,2,3-cd)pyrene, naphthalene | In-house modified STN EN ISO 6468 |
| GC-Flame retardants | brominated diphenylethers, hexabromocyclododecane (HBCDD) | In house modified EPA 1614 GC-MS/MS, GC-MS(NCI) method |
| Metals | cadmium and its compounds, lead and its compounds, mercury and its compounds, nickel and its compounds | ISO 17294-2 - ICP-MS |
| GC - Industrial | chloroalkanes, C10-13 | In house GC-MS(NCI) method |
| GC - Pesticides | aldrin, dieldrin, endrin, isodrin, DDT total, dichlorvos, dicofol, endosulfan, heptachlor and heptachlor epoxide, hexachlorobenzene, hexachlorocyclohexane (hexachlorocyclohexane all isomers), para-para-DDT, pentachlorobenzene, chlorfenvinphos, chlorpyrifos (chlorpyrifos ethyl), trifluralin | In-house modified STN EN ISO 6468 |
| GC - Industrial | di(2-ethylhexyl)phthalate (DEHP) | In-house modified STN EN ISO 6468 |
| GC – Industrial / Pesticide | nonylphenols (4-nonylphenol), octylphenols ((4-(1,1’,3, 3’-tetramethyl-butyl)-phenol)), pentachlorophenol | In-house modified STN EN ISO 6468 |
| LC - Industrial | perfluoroctane sulfonic acid and its derivatives (PFOS) | In-house modified STN EN ISO 6468 |
| GC – VOCs | benzene, 1,2-dichloroethane, dichloromethane, hexachlorobutadiene, trichlorobenzenes, trichloromethane (chloroform), carbon-tetrachloride, tetrachloro-ethylene, trichloro- ethylene | In-house GC-MS/MS method in conjunction with static headspace |
| GC - Biocides | tributyltin compounds (tributyltin-cation) | In-house method, 50 ml of raw water sample - SBSE-GC-MS/MS - high sensitivity, derivatization needed |

Five main methods were used to determine selected Priority substances: GC-MS/MS, GC-MS(NCI), LC-HRMS, LC- MS/MS and ICP-MS.

1. GC-MS/MS analyses were conducted by Agilent 7010 Triple Quadrupole (Agilent Technologies Deutschland, Waldbronn, Germany).

An DB-5mc column (30 m x 0.25 mm, 0.25 μm) from Agilent (Germany) coupled to the back-flush system was used at a constant flow of 1.6 ml/min. Hydrogen was used as mobile phase. The temperature gradient programme was initialized at 50 °C, hold for 2 min, increased to 130 °C with gradient 18 °C/min, increased to 290 °C with gradient 7 °C/min, increased to 320 °C (hold 5 min) with gradient 18 °C/min. The sample introduction used stir-bar sorptive extraction (SBSE), where the stir bar covered with PDMS stationary phase was placed into the liner of the injector and subsequently the trapped substances were thermally released into the separation system. Data acquisition was performed in multiple reaction monitoring (MRM) mode. This system was also equipped with the static headspace option for analyzing VOC compounds.

2. GC-MS(NCI) analyses were conducted by a Agilent 5975C Triple axis Quadrupole (Agilent Technologies Deutschland, Waldbronn, Germany) working in the Negative Chemical Ionization (NCI) mode with methane.

An DB-5mc column (15m x 0.25 mm, 0.25 μm) from Agilent (Germany) coupled to the back-flush system was used at a constant flow of 2.0 ml/min. Hydrogen was used as mobile phase. The temperature gradient programme was initialized at 50°C, held for 2 min, increased to 320°C (hold 5 min) with gradient 20 °C/min. The sample introduction used PTV injector allowing to inject up to 50 µl of sample extract. Data acquisition was performed in selected ion monitoring (SIM) mode. This system was used for compounds having electronegative atoms.

3. LC-HRMS (LC-ESI-QTOF-MS) analyses were conducted using an ultra high-performance liquid chromatographic (UHPLC) system with a HPG-3400 pump (Dionex Ultimate 3000 RSLC, Thermo Fischer Scientific, Dreieich, Germany) coupled to a QTOF mass spectrometer (Maxis Impact, Bruker Daltonics, Bremen, Germany).

Chromatographic separation was performed using an Acclaim RSLC C18 column (2.1 x 100 mm, 2.2 μm) from Thermo Fischer Scientific (Dreieich, Germany) preceded by a guard column of the same packaging material, thermostated at 30 °C. For positive ionization mode (PI), the mobile phases were water/methanol 90/10 (solvent A) and methanol (solvent B) both containing 5 mM ammonium formate and 0.01 % formic acid. For negative ionization mode (NI), the mobile phases consisted of water/methanol (solvent A) and methanol (solvent B) both acidified with 5 mM ammonium acetate. The adopted gradient elution programme was the same for both ionization modes, starting with 1% B with a flow rate of 0.2 ml min-1 for 1 min and it increased to 39% in 2 min (flow rate 0.2 mL min-1), and then to 99.9 % (flow rate 0.4 ml min-1) in the following 11 min. Then, it was kept constant for 2 min (flow rate 0.48 ml min-1), then the initial conditions were restored within 0.1 min, kept for 3 min and then the flow rate decreased to 0.2 ml min-1. The injection volume was 5 μl.

The quadrupole time-of-flight mass spectrometry (QTOF-MS) system was equipped with an electrospray ionization interface (ESI), operating in positive and negative mode, with the following operation parameters: capillary voltage 2500 V (PI) and 3500 V (NI); end plate offset, 500 V; nebulizer pressure 2 bar; drying gas 8 l min−1 and gas temperature 200°C. The QTOF-MS system was operated in data-independent (broadband collision-induced dissociation (bbCID)) acquisition, as well as in data dependent (Auto MS/MS) acquisition mode and recorded spectra over the range m/z 50−1200, with a scan rate of 2 Hz. A QTOF-MS external calibration was performed daily with the manufacturer's solution. The instrument provided a typical resolving power (FWHM) between 36,000-40,000 at m/z 226.1593, 430.9137 and 702.8636.

4. LC-ESI-MS/MS analyses were conducted by a Thermo UHPLC Accela system interfaced with a Thermo TSQ Quantum Access Triple Quadrupole Instrument (Thermo, San Jose, CA, USA).

An X-Bridge C18 column (2.1 x 50 mm, 2.5 μm) from WATERS (Dublin, Ireland) preceded by a guard column was used at a constant flow of 100 μl min−1, thermostated at 30°C. The mobile phases consisted of water containing 0.1 % formic acid (v/v) (solvent A) and methanol (solvent B). The gradient programme was initialized with 50% B, increased to 100% B in 9 min and then it kept constant for 3 min. Afterwards, the initial conditions were restored within 1.5 min and kept constant for 3.5 min. The injection volume was set up to 5 µl. The ESI probe was operated in Positive Ionization. ESI voltage was set at 4000 V. The sheath gas (N2) flow rate was set at 40 A.U (Arbitrary Units), the auxiliary gas (N2) flow rate was set at 20 A.U and the ion transfer capillary temperature was set at 270 ⁰C. Data acquisition was performed in multiple reaction monitoring (MRM) mode.

5. ICP-MS analyses were conducted by Agilent 7900 ICP-MS with standard nickel cones. Agilent ICP-MS systems include four argon gas lines as standard. The optimum nebulizer gas flow was 0.7 l/min, so an additional 0.5 l/min make-up flow was used to give the optimum total carrier flow. All solutions (diluent, carrier, rinse, internal standard, calibration standards) were prepared using 2% HNO3 and 1% HCl. HCl was included to ensure the stability of Hg in solution.

### Quality Assurance & Quality Control (QA/QC)

Quality assurance and quality control (QA/QC) was applied during the sample preparation and instrumental analysis. A mix of internal standards was added into each sample prior to extraction to ensure satisfactory recovery of the target compounds. Samples spiked with a mix of known contaminants were also analyzed in each batch of samples. Moreover, simultaneous process and monitoring of solvent blanks and procedural blanks during the analysis of all samples was realized to assess any external contamination which might have been brought in during the sample preparation of the extracts and analysis. A mix of known analytes (RTI calibrant substances) was used to assess the stability of retention time during instrumental analysis. A QC sample, including a representative number of contaminants included in the target list, was injected every five injections for the evaluation of good operation and high sensitivity of the instrumentation.

# Results

## In situ measured parameters

Annex 2 contains an overview of the parameters that were measured in the field, briefly characterised as follows.

- Dissolved oxygen, oxygen saturation

  - The oxygen levels measured at sites Otaci and Nimereuca in April 2025 were quite low for Dniester River. At Otaci even lower than the minimum in the routine 2009 - 2024 surface water quality monitoring data.

  - Substantially lower oxygen levels at several tributary sampling sites indicate impacts of wastewater discharges.

  - Oxygen levels at Botna-Chircăiești though indicated oversaturation in July and October, significantly higher than regular oxygen saturation levels in rivers (range 80 - 120%). Oversaturation was also the case in July at Nimereuca.

- pH

  - Noteworthy mentioning is pH 4.7 measured at Otaci in April. This is not a value one can expect in a river like the Dniester. However, in April the samples were not taken from the bridge, but from a small inlet further downstream at the right bank.

- Hardness

  - Hardness is relevant for interpreting the cadmium concentrations; refer to Annex 1, entry (6) Cadmium and its compounds (depending on water hardness classes) for further details. The EQSs of Class 5: ≥ 200 mg CaCO3/l have been applied for the assessments in this report.

- Conductivity

  - The conductivity at the three Dniester River sites is considerably lower than at the tributary sites. The latter can be caused by a combination of factors, such as discharges of wastewater and/or salinisation due to irrigation. The Moldovan tributaries are quite small in terms of size and flow, therewith more prone to such factors than the Dniester River. Also geogenic characteristics might play a role.

The sampling team noticed both visually and by smell that the water sampled downstream Bălţi and Goian, as well as at Gura Bîcului was substantially affected by wastewater.

In July 2025, the water level at most tributary sites was low and at Chircăiești the Botna River was not flowing at all. Low water levels were also noticed in October 2025 at the sites Ustia, Goian and Chircăiești.

## Priority substances not detected

34  - individual - Priority substances were not detected in all water samples (Table 3-1).

Table 3-1 Individual Priority substances not detected (<LOD)

| Group | Substance name | Unit |  | LOD |
| --- | --- | --- | --- | --- |
| OCP | Aldrin | ng/l |  | 0.04 |
|  | Dichlorvos | ng/l |  | 0.04 |
|  | Dicofol | ng/l |  | 0.04 |
|  | Endosulfan I | ng/l |  | 0.04 |
|  | Endosulfan II | ng/l |  | 0.04 |
|  | Endrin | ng/l |  | 0.04 |
|  | Heptachlor | ng/l |  | 0.04 |
|  | Heptachlor epoxide B | ng/l |  | 0.004 |
|  | Isodrin | ng/l |  | 0.04 |
|  | p,p'-DDT | ng/l |  | 0.04 |
|  | Pentachlorobenzene | ng/l |  | 0.04 |
|  | Pentachlorophenol | ng/l |  | 20 |
| PEST | Aclonifen | ng/l |  | 0.61 |
|  | Alachlor | ng/l |  | 0.11 |
|  | Bifenox | ng/l |  | 0.65 |
|  | Cybutryne | ng/l |  | 0.94 |
|  | Cypermethrin | ng/l |  | 0.16 |
|  | Quinoxyfen | ng/l |  | 0.59 |
|  | Trifluralin | ng/l |  | 0.04 |
| PAH | Benzo(a)pyrene | ng/l |  | 0.03 |
|  | Benzo[k]fluoranthene | ng/l |  | 0.03 |
| VOC | 1,2-dichloroethane | µg/l |  | 0.32 |
|  | Benzene | µg/l |  | 0.04 |
|  | Carbon-tetrachloride | µg/l |  | 0.035 |
|  | Dichloromethane | µg/l |  | 0.33 |
|  | Hexachlorobutadiene | µg/l |  | 0.032 |
|  | Tetrachloro- ethylene | µg/l |  | 0.1 |
|  | Trichloro- ethylene | µg/l |  | 0.15 |
|  | Trichlorobenzenes | µg/l |  | 0.03 |
| IND | 4-Nonylphenol | µg/l |  | 0.02 |
|  | 4-Octylphenol | µg/l |  | 0.02 |
|  | Brominated diphenylethers | ng/l |  | 0.03 |
|  | Hexabromocyclododecane (HBCDD) | ng/l |  | 0.03 |
|  | Octylphenols ((4-(1,1’,3, 3’-tetramethyl-butyl)-phenol)) | µg/l |  | 0.02 |

However, the AA-EQS concentrations of the substances listed in Table 3-2 below are lower than the limit of detection (LOD). Which means that no significant conclusions can be drawn from <LOD entries. The substance might be present in the water in concentrations exceeding the AA-EQS.

Table 3-2 Substances with AA-EQS lower than LOD

| Substance | Unit | LOD | AA-EQS1 | MAC-EQS |
| --- | --- | --- | --- | --- |
| Heptachlor | ng/l | 0.04 | 0.0002 | 0.3 |
| Heptachlor epoxide B | ng/l | 0.004 | 0.0002 | 0.3 |
| Cypermethrin | ng/l | 0.16 | 0.08 | 0.6 |

## Detected Priority substances

Table 2-1 below provides an overview of the Priority substances that have been detected in one or more samples taken in April, July and/or October 2025. In total 28 individual Priority substances were detected.

The tables in Annex 5 contain overviews of all sites for each sampling month.

Table 3-3 Overview of substances detected in one or more samples taken in April, July and October 2025

|  |  |  | Dniester |  |  |  | Răut |  |  | Ichel |  | Bîc |  |  | Botna |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Group | Substance name | Unit | Otaci | Nimereuca | Palanca |  | Bălţi | Ustia |  | Goian |  | Străşeni | Gura Bîcului |  | Chircăiești |
| HM | Cadmium and its compounds | µg/l |  |  |  |  | √ |  |  |  |  |  |  |  |  |
|  | Lead and its compounds | µg/l | √ | √ | √ |  | √ | √ |  | √ |  | √ | √ |  | √ |
|  | Mercury and its compounds | µg/l |  |  |  |  |  |  |  | √ |  |  | √ |  | √ |
|  | Nickel and its compounds | µg/l | √ | √ | √ |  | √ | √ |  | √ |  | √ | √ |  | √ |
| OCP | 1,2,3,4,5,6-Hexachlorocyclohexane (HCH total) | ng/l | √ | √ | √ |  | √ | √ |  | √ |  |  | √ |  | √ |
|  | Chlorfenvinphos | ng/l |  |  |  |  | √ |  |  |  |  |  | √ |  | √ |
|  | Chlorpyrifos (Chlorpyrifos-ethyl) | ng/l |  |  |  |  | √ |  |  | √ |  |  | √ |  |  |
|  | DDT total | ng/l |  |  |  |  | √ |  |  | √ |  |  | √ |  | √ |
|  | Dieldrin | ng/l |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
|  | Heptachlor epoxide A | ng/l |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
|  | Hexachlorobenzene | ng/l |  |  |  |  | √ | √ |  |  |  |  |  |  |  |
| PEST | Atrazine | ng/l | √ | √ | √ |  | √ | √ |  | √ |  | √ | √ |  | √ |
|  | Diuron | ng/l |  |  | √ |  | √ | √ |  | √ |  | √ | √ |  | √ |
|  | Isoproturon | ng/l |  |  |  |  |  |  |  |  |  |  | √ |  |  |
|  | Simazine | ng/l |  |  |  |  |  | √ |  | √ |  | √ |  |  |  |
|  | Terbutryn | ng/l | √ | √ | √ |  | √ | √ |  | √ |  | √ | √ |  | √ |
| PAH | Anthracene | ng/l | √ | √ |  |  | √ | √ |  | √ |  |  | √ |  |  |
|  | Benzo(b)fluoranthene | ng/l |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
|  | Benzo[g,h,i]perylene | ng/l |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
|  | Fluoranthene | ng/l | √ | √ |  |  | √ |  |  | √ |  |  | √ |  |  |
|  | Indeno(1,2,3-cd)pyrene | ng/l |  |  |  |  |  |  |  |  |  |  | √ |  |  |
| VOC | Naphtalene | µg/l | √ |  |  |  | √ |  |  |  |  |  |  |  |  |
|  | Trichloromethane (chloroform) | µg/l |  |  |  |  | √ | √ |  | √ |  |  | √ |  |  |
| IND | Chloroalkanes, C10-13 | ng/l |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
|  | Di(2-ethylhexyl)phthalate (DEHP) | µg/l | √ | √ | √ |  | √ |  |  | √ |  | √ | √ |  | √ |
|  | Nonylphenols | µg/l |  |  | √ |  | √ | √ |  | √ |  |  | √ |  |  |
| TBT | Tributyltin compounds | ng/l | √ | √ | √ |  |  |  |  |  |  |  |  |  |  |
| PFOS | Total Perfluorooctanesulfonic acid (PFOS) | ng/l | √ | √ | √ |  | √ | √ |  | √ |  | √ | √ |  | √ |
| Total № of detected substances |  |  | 11 | 10 | 10 |  | 23 | 12 |  | 16 |  | 8 | 23 |  | 11 |

### AA-EQS, MAC-EQS

The Directive 2013/39/EU contains two sets of environmental quality standards (EQSs) for assessment of the chemical status:

- AA-EQS: annual average environmental quality standard

  - This EQS is to be applied to the arithmetic mean of the number of samples taken throughout the year. WFD Annex V.1.3.4 indicates monthly sampling, thus 12 samples per year.

- MAC-EQS: maximum allowable concentration environmental quality standard

  - This is a figure not to be exceeded any time, in any of the samples taken throughout the year.

In case concentrations of all (detected) Priority substances comply with their AA-EQS and MAC-EQS, then a surface waterbody can be qualified as being of ‘Good chemical status’.

Otherwise, the waterbody has to be qualified as ‘Failing to achieve good chemical status’.

Referring to the AA-EQS in the subsections below is merely for providing a dimension to the results, thus not intended as compliance checking (refer also to subsection 4.2.6 in Chapter 4). The three survey samples are too low a number for statistically relevant checking compliance with the AA-EQS.

### Heavy metals

The table below and other such ‘screenshot tables’ in the following subsections are merely provided to give a quick impression. Detailed tables are included in Annex 5.

Table 3-4 Overview of detected heavy metals (screenshot)

Note that the concentrations of the metals were determined in the dissolved fraction (after filtration over a 0.45 µm mesh filter).

Cadmium was detected only in the sample taken downstream Bălţi in July 2025, with a concentration near the LOD.

Lead was detected at all locations, although not in all months, with concentrations below the AA-EQS.

Nickel was also detected at all locations and not in all months. However, nickel concentrations were higher than the AA-EQS in three cases: Ichel-Goian (October), Bîc-Străşeni (October), and Botna-Chircăiești (July).

Mercury was detected in samples taken in April 2025. The concentration of the sample taken at Ichel-Goian was 0.011 µg/l, rather close to the LOD (0.006 µg/l). The concentrations in the samples taken at Bîc-Gura Bîcului and Botna-Chircăiești exceed the MAC-EQS.

### Organochlorine pesticides

Hexachlorocyclohexanes were detected in the samples taken at the three sites along the Dniester River and in the tributaries, with concentrations well below the AA-EQS.

The other detected organochlorine pesticides (OCPs) were only found in tributary samples.

Table 3-5 Overview of detected OCPs (screenshot)

Table 3-5 shows several relevant details, like:

- Presence/absence between sampling sites and between sampling months.

- The differences between the two sites along the Bîc River: no OCPs detected at Străşeni, but several ones at Gura Bîcului.

- The presence of DDT, DDD and DDE, notably in the samples taken at Chircăiești.

- Dieldrin being detected downstream Bălţi and at Gura Bîcului.

Noteworthy mentioning are the concentrations of chlorpyrifos in the samples taken in the Răut River downstream Bălţi. With the concentration in the April sample being higher than the AA-EQS and in July exceeding the MAC-EQS.

Heptachlor epoxide A was detected in the July samples taken downstream Bălţi and at Gura Bîcului, with concentrations higher than the AA-EQS. However, this is a good example of results that should be interpreted with care:

- The AA-EQS – 0.0002 ng/l – is lower than the LOD (0.004 ng/l).

- The reported concentrations, 0.005 respectively 0.006 ng/l, are between the LOD (0.004 ng/l) and the LOQ (0.012 ng/l), a range within which results of laboratory analysis are less accurate; refer to the discussion in the next chapter (subsection 4.2.6).

### Other pesticides

Five more pesticides were detected (Table 3-6).

Table 3-6 Other detected pesticides (screenshot)

Isoproturon was detected in the October sample taken at Gura Bîcului only.

All concentrations were below their related AA-EQS, although the diuron concentration of the sample taken in April at Gura Bîcului – 192 ng/l – was near to the AA-EQS (200 ng/l).

### Polyaromatic hydrocarbons

Five PAHs were detected: anthracene, benzo(b)fluoranthene, benzo[g,h,i]perylene, fluoranthene, and indeno(1,2,3-cd)pyrene, in various combinations between the sites and/or months. None of these PAHs were detected at Dniester-Palanca, Bîc-Străşeni and Botna-Chircăiești.

Table 3-7 Overview of detected PAHs (screenshot)

The anthracene and fluoranthene concentrations were below the related AA-EQSs. However, the fluoranthene concentrations of the samples taken Răut-Bălţi in April (4.6 ng/l) and July (5.2 ng/l) were nearing the AA-EQS (6.3 ng/l).

### Volatile organic compounds

Table 3-8 Overview of detected VOCs (screenshot)

Three VOCs were detected, while noticing that toluene is not included in the Directive 2013/39/EU.

The naphthalene and toluene concentrations in the April sample taken at Otaci are remarkable, especially within the overall picture. Remarkable, because:

- Volatile organic compounds are compounds that have a high vapour pressure and low water solubility.

  - So, even when present in the river, VOCs would evaporate rather quickly.

- Naphthalene was detected at Otaci also in July, although with a concentration just above the LOD (0.004 µg/).

  - The only other occurrence of naphthalene was in the July sample taken downstream Bălţi, also with a concentration just above the LOD.

- Toluene was detected in July at Palanca just above the LOD and in more significant concentrations in the April and October samples taken downstream Bălţi.

- Emissions of naphthalene and toluene are normally due to human activities.

  - However, they were e.g. not detected at all downstream Chisinau at Gura Bîcului. There is no agglomeration of such size upstream between Otaci and Naslavcea, near the upstream border.

In April, the sample at Otaci was not taken from the bridge, due to ‘force majeure’. The sample was taken from a small inlet at the right bank further downstream. As mentioned in section 3.1, the pH was extremely low: 4.7. Together with the abovementioned notices, the April 2025  findings for naphthalene (and toluene) are considered questionable. Definitely in terms of being representative for the overall water quality of the Dniester River at Otaci as such.

Trichloromethane (chloroform) was detected in all three samples taken at Răut-Bălţi and at Bîc- Gura Bîcului. Furthermore it was detected in the October sample taken at Răut-Ustia and in the July sample taken at Ichel-Goian. All concentrations were below the AA-EQS.

### Industrial pollutants

The detected Priority substances labelled ‘Industrial pollutants’ are listed in Table 2-1.

Table 3-9 Detected ‘Industrial pollutants’ (screenshot)

However, phthalates are an indicator for ‘plastic’ and can originate from various sources, including domestic solid waste dumped into rivers. Which can explain that di(2-ethylhexyl)phthalate was detected at most sites, whereas chloroalkanes were solely found downstream Bălţi and at Gura Bîcului.

Nonylphenols were also detected downstream Bălţi and at Gura Bîcului, as well as in the April sample taken downstream Goian. The concentration of the sample taken at Răut-Bălţi in July was higher than the AA-EQS.

### Tributyltin compounds

Tributyltin compounds (TBT) were only found in the samples taken from the Dniester River, with concentrations higher than the AA-EQS and in several samples even exceeding the MAC-EQS.

Table 3-10 Concentrations of tributyltin compounds (screenshot)

The representativeness of the site where the samples at Otaci were taken in April 2025 is considered questionable (subsection 3.3.6 above). However, within the overall picture there can be virtually no doubt about the presence of tributyltin compounds in the Dniester River as a whole at Otaci in April 2025.

### Perfluorinated compounds, PFOS

Table 3-11 Detected perfluorinated compounds (screenshot)

PFOS were found at all sites in April and July, but not at all in the October samples.

The total PFOS concentrations were higher than the AA-EQS (0.65 ng/l) in July samples taken downstream Bălţi and at Gura Bîcului, but very well below the MAC-EQS OF 36,000 ng/l.

### Per- and poly-fluorinated alkyl substances, PFAS

The Environmental Institute analysed 38 PFAS derivatives, which are not included in the Directive 2013/39/EU. However, PFAS are included in the proposal for a Directive amending Directive 2000/60/EC establishing a framework for Community action in the field of water policy, Directive 2006/118/EC on the protection of groundwater against pollution and deterioration and Directive 2008/105/EC on environmental quality standards in the field of water policy.

Details about the results of analysis of PFAS are included in Annex 6.

18 PFAS derivatives were not detected in any water sample. The remaining 20 derivatives were detected at all sites, albeit in various combinations.

For orientation purposes: in the proposal the MAC-EQS for “Per- and poly-fluorinated alkyl substances (PFAS) – sum of 24” is set to “Sum of PFOA equivalents 0.0044 µg/l” (=4.4 ng/l).

# Discussion

Sixty two (62) individual Priority substances mentioned in the Directive 2013/39/EU have been analysed under this Research monitoring programme.

WFD Annex V. 1.3.4 indicates monthly sampling for Priority samples, thus 12 samples per year. There are already quite some differences between the 3 samples taken at each site, both in terms of presence/absence, as well as concentrations.

This chapter is dedicated to how the results might be applied to a wider scale. Please notice that there is room for interpretation!

## Priority substances not detected

Thirty four (34) individual Priority substances were not detected in any of the samples taken at the nine sites in April, July and October 2025.

### Sample media

It is important to notice that ‘water’ is not always the preferred media for monitoring & assessment of Priority substances.

Some Priority substances are ‘hydrophobic’, meaning that they poorly dissolve in water and are rather absorbed to suspended solids and sediment. Substances can furthermore accumulate in organisms like mussels and fish, implying ‘biota’ (tissue samples) as preferred sample medium.

Suspended solids, sediment and biota could not be included in the Research monitoring programme, due to logistical and financial implications.

### LOD > AA-EQS

In three instances the limit of detection was higher than the AA-EQS, implying the substance might be present in the water in concentrations exceeding the AA-EQS.

Table 4-1 Cases with LOD > AA-EQS

| Group | Compound | Unit | LOD | AA EQS | MAC EQS |
| --- | --- | --- | --- | --- | --- |
| OCP | Heptachlor | ng/l | 0.04 | 0.0002 | 0.3 |
|  | Heptachlor epoxide B | ng/l | 0.004 | 0.0002 | 0.3 |
| PEST | Cypermethrin | ng/l | 0.16 | 0.08 | 0.6 |

### Extrapolation

With not having detected substances in water samples taken at 9 sites in 3 different seasons, it reasonably can be assumed that they might not occur in the water-phase in most of the surface waterbodies inside the Moldovan Dniester River Basin Management District.

Bălţi and Chișinău are the largest cities of Moldova, with plenty small and medium-sized enterprises. When substances are not detected downstream these cities, then why to expect them downstream other settlements, unless there are certain specific industrial activities?

This, while acknowledging that the wastewaters of Bălţi and Chișinău are treated in wastewater treatment plants. Notably hydrophobic pollutants can be reduced throughout the various stages of wastewater treatment.

Pesticides are commonly associated with agricultural activities, its application also depending on the types of crops and their growing seasons. From this point of view one can expect regional differences. The results for organochlorine pesticides between the samples taken at Straseni (none detected) and at Guru Bîcului show substantial differences even within one tributary subbasin (subsection 3.3.3).

Nevertheless, the tributary sampling sites at least give impressions from north to south. So why to expect certain pesticides in specific – other – parts of the Dniester RBMD? Such questions not necessarily always have to be solved via more sampling and laboratory analysis. Desk studies (e.g. inventory of pesticide sales, research articles, et cetera) also can reveal potential ‘hot spots’. Of course, laboratory analysis of field samples is required for checking whether or not the suspects are indeed in the water, sediment and/or biota.

However, with overall resources for monitoring and laboratory analysis being limited, it prevails for the short to medium term to focus on those Priority substances that have been detected.

## Priority substances detected

It is good to notice that the samples are not related in time. For example, the samples at Otaci and Nimereuca were taken within a few hours, whereas it can take a couple of days for the Dniester River water to travel from Otaci to Nimereuca.

### LOD, LOQ

AA-EQS, MAC-EQS The Directive 2013/39/EU contains two sets of environmental quality standards for assessment of the chemical status - AA-EQS: annual average environmental quality standard - This EQS is to be applied to the arithmetic average/mean of a number of samples taken throughout the year. WFD Annex V. 1.3.4 indicates monthly sampling, thus 12 samples per year. - MAC-EQS: maximum allowable concentration environmental quality standard - This is a figure not to be exceeded any time, thus in none of the samples. In case the concentrations of all (detected) Priority substances comply with the AA-EQS and MAC-EQS, then a surface waterbody can be qualified as being of ‘Good’ chemical status. Otherwise, the waterbody has to be qualified as ‘Failing to achieve good’ chemical status.

Comparison of the concentrations of the detected Priority substances with the AA-EQS and MAC-EQS of the Directive 2013/39/EU provide a certain perspective, but one should remain cautious doing so.

The Directive 2009/90/EC  mentions the following in its Article 4 Minimum performance criteria for methods of analysis: “1. Member States shall ensure that the minimum performance criteria for all methods of analysis applied are … a limit of quantification equal or below a value of 30 % of the relevant environmental quality standards.”

The Environmental Institute calculated the limit of quantification (LOQ) as three times the limit of detection (LOD). Implying that for comparison with the EQSs the LOD should be better than EQS*0.3/3 = 10% of the EQS. Refer to Annex 3 for such details.

In three cases of detected substances the LOD was ≥10% of EQS (Table 4-2).

Table 4-2 Detected substances with LOD ≥ 10% of EQS

| Group | Compound | Unit | LOD | LOQ | AA- EQS | MAC- EQS |
| --- | --- | --- | --- | --- | --- | --- |
| HM | cadmium and its compounds | µg/l | 0.03 | 0.10 | 0.25 | 1.5 |
| OCP | heptachlor epoxide A | ng/l | 0.004 | 0.01 | 0.0002 | 0.3 |
| TBT | tributyltin compounds | ng/l | 0.04 | 0.1 | 0.2 | 1 |

Heptachlor epoxide A illustrates the case very well:

- Not having detected heptachlor epoxide A doesn’t imply that it might not be there. It might be present in concentrations above the AA-EQS, but the LOD of laboratory analysis was simply not sufficiently sensitive with best available technology.

- It was detected in the July samples taken downstream Bălţi and at Guru Bîcului, with concentrations of respectively 0.005 ng/l and 0.006 ng/l. Just above the LOD, a concentration range statistically not being robust.

### Substances detected at ‘critical levels’: MAC-EQS

Just one sample with a concentration exceeding the MAC-EQS by definition results in the ‘Failing to achieve good’ chemical status.

This was the case with three substances:

- Mercury at Bic-Gura Bîcului and Botna-Chircăiești (April 2025 samples).

- Chlorpyrifos at the site Răut-Bălţi (July 2025 sample) .

- Tributyltin compounds at the Dniester River sites Otaci and Nimereuca (several instances).

Their LOQ was well below 30% of the MAC-EQS.

### Substances detected at ‘near critical levels’: AA-EQS

WFD Annex V.1.3.4 indicates a monthly sampling frequency, 12 samples per year, for Priority substances. The three samples taken per site during the Research monitoring programme already showed quite some variations. Thus, three survey samples with concentrations higher or near their AA-EQSs should not be used in terms of compliance checking. Nevertheless, results around the AA-EQS instigate attention.

As with all boundaries, there is a certain grey zone. For example, the diuron concentration of the April sample taken at Gura Bîcului was 192 ng/l, whereas the AA-EQS is 200 ng/l; not that significant a difference. Therefore, the following criterion was added for identifying Priority substances of interest:

- Concentrations ≥ 0.66 * AA-EQS also might indicate certain risks of not achieving good chemical status.

### Not detected at all sites

This may sound like a trivial observation but is – inter alia –  relevant for the design of surface water quality monitoring programmes. And, of course, from a water management perspective.

Take, for example, the results of the samples taken along the Bîc River at Străşeni and at Gura Bîcului. Several Priority substances were detected at Gura Bîcului that were not found in the samples taken at Străşeni, including pesticides.

Also notice differences between substances (not) detected at Ichel-Goian versus at Bîc-Străşeni; Goian and Străşeni are both relatively small towns.

More organochlorine pesticides were detected more frequently in the samples taken at Botna-Chircăiești when compared with e.g. Bîc Străşeni and Ichel-Goian (Table 3-5).

The results of the tributyltin compounds are perhaps most straightforward: they were detected only in the Dniester River samples from Otaci onwards, with significant concentrations.

### Not at all times

With pesticides one can expect seasonal variations in concentrations.

Mercury as such is though not necessarily associated with pesticides, but was only found in the April samples taken at Ichel-Goian, Bîc-Gura Bîcului and Botna-Chircăiești. The concentrations in the latter two samples exceeded the MAC-EQS, which – by definition – means ‘Failing to achieve good’ chemical status’.

The hydrological regime of Moldovan tributaries is an important factor to be taken into account. At various sites the sampling team observed little to no flow in July; compare for example Figure 4-1 below.

One might argue that emissions of mercury from a (unknown) source upstream simply did not reach Chircăiești, since the Botna River didn’t flow. Relevant a notice when comparing data with the EQSs.

Figure 4-1 Water levels at the Botna- Chircăiești sampling sites in April, July 2025

### Hydrological considerations: Bălţi – Ustia

Less substances were detected at the sampling site Ustia in comparison with the samples taken downstream Bălţi.

Possible reasons for having detected less substances at Ustia could be:

- Sedimentation of hydrophobic substances absorbed to suspended solids.

- Dilution. The average flow of  the Răut River downstream at Jeloboc is 5.6 times larger than at Bălţi (compare Table 4-3). Implying that at Ustia concentrations could be about 20% of the concentrations at Bălţi, due to dilution.

The above considerations were not pursued for this report. However, in combination with the previous subsection (no flow at Chircăiești during sampling in July 2025), they illustrate that hydrological aspects are relevant for interpretation of surface water quality data.

Table 4-3 Average, minim and maximum flows of Răut River at Bălţi and Jeloboc

## Synthesis

This chapter started with mentioning “Please notice that there is room for interpretation!”; readers might weigh the raw data and considerations differently. If only a comprehensive overview of ‘the chemical status’ of all delineated surface waterbodies inside the Moldovan Dniester RBMD could be assessed merely by taking samples at nine sites, three times in a year…

The ‘Priority substances and certain other pollutants’ considered of short/medium term attention are summarised in the table below.

Table 4-4 Detected Priority substances considered of short/medium term attention

|  | DNIESTER |  |  |  | RĂUT |  |  | ICHEL |  | BÎC |  |  | BOTNA |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Substance | Otaci | Nimereuca | Palanca |  | Bălţi | Ustia |  | Goian |  | Străşeni | Gura Bîcului |  | Chircăiești |
| Mercury |  |  |  |  |  |  |  |  |  |  | √ |  | √ |
| Chlorpyrifos |  |  |  |  | √ |  |  |  |  |  |  |  |  |
| Tributyltin compounds | √ | √ | √ |  |  |  |  |  |  |  |  |  |  |
| Nickel |  |  |  |  |  |  |  | √ |  | √ | √ |  | √ |
| Heptachlor epoxide A (*) |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
| PFOS |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
| Nonylphenols |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
| Chlorfenvinphos |  |  |  |  |  |  |  |  |  |  |  |  | √ |
| Dieldrin |  |  |  |  |  |  |  |  |  |  | √ |  |  |
| Diuron |  |  |  |  |  |  |  |  |  |  | √ |  |  |
| Terbutryn |  |  |  |  |  |  |  |  |  |  | √ |  |  |
| Fluoranthene |  |  |  |  | √ |  |  |  |  |  |  |  |  |

(*) Concentrations were just above the LOD, hence results should be interpreted with caution.

| √ | one or more concentrations >MAC-EQS |
| --- | --- |
| √ | one or more concentrations >AA-EQS |
| √ | one or more concentrations >0.66 * AA-EQS |

Requiring attention in two ways:

- For further capacity building of the Water Quality Laboratory.

- For design of the surface water quality monitoring programmes.

### Capacity development of the Water Quality Laboratory

The Water Quality Laboratory under the Environmental Reference Laboratory of the Environment Agency analysed the following selected Priority substances within the routine surface water quality monitoring programmes throughout the year 2024.

Table 4-5 Selected Priority substances analysed by the Water Quality Laboratory

| Substance | LOD | AA-EQS | MAC-EQS |
| --- | --- | --- | --- |
| Mercury | 0.03 µg/l | - | 0.07 µg/l |
| Nickel | 0.165 µg/l | 4 µg/l | 34 µg/l |
| Heptachlor epoxide A | 7 ng/l | 0.0002 ng/l | 0.3 ng/l |
| Dieldrin | 9 ng/l | ∑ drins =10 ng/l | - |
| Fluoranthene | 7.5 ng/l | 6.3 ng/l | 120 ng/l |

- The LOD for heptachlor epoxide A is higher than the AA-EQS and the MAC-EQS. Obviously, the LOD has to be lowered in order to obtain relevant data, although this will be technically challenging, notably regarding the AA-EQS.

- The LOD for fluoranthene is higher than the AA-EQS, implying that no statistically reliable data can be collected for comparison with the AA-EQS.

- The LOD of dieldrin is just below the AA-EQS for the sum of drins (aldrin, dieldrin, endrin, isodrin), therewith inadequate for compliance checking.

- The LOD of mercury is higher than the required 10% of the MAC-EQS, introducing less robust compliance checking.

- The LOD of nickel is sufficiently low.

From the above one can infer that an approach could be ranked as follows.

- Optimise analysis of selected Priority substances that are currently monitored.

  - The laboratory should investigate whether optimisation can be achieved with the currently available equipment. Options might involve for example sample pretreatment and/or using other (EN/ISO) methods of analysis.

- Investigate which of the other selected Priority substances (Table 4-4) might be analysed with the currently available equipment.

  - For example taking into consideration recommended methods of analysis, necessary reagents and consumables, et  cetera.

- Develop a strategy for enhancing the capacity for analysis requiring new analytical equipment, implying also new methods of analysis, reagents and consumables, et cetera.

  - Here it would be proper to anticipate the proposal for a Directive amending Directive 2000/60/EC establishing a framework for Community action in the field of water policy, Directive 2006/118/EC on the protection of groundwater against pollution and deterioration and Directive 2008/105/EC on environmental quality standards in the field of water policy.

Receiving training and obtaining accreditation by the National Accreditation Centre of the Republic of Moldova (MOLDAC) apply in all scenarios.

### Design of surface water quality monitoring programmes

Table 4-4 indicates that monitoring of the selected Priority substances can be limited to certain locations or stretches of tributaries.

For example, tributyltin compounds were detected only at the Dniester River sampling sites. There is no doubt that the source(s) of tributyltin compounds’ emissions are situated upstream Otaci. Monitoring tributyltin compounds at Otaci could suffice for tracking trends over time.

The results for the other selected Priority substances though merely indicate emission sources further upstream the sampling site that are yet to be pinpointed.

Typical cases for ‘investigative monitoring’. Preferably preceded by/combined with desktop studies. For example, for ‘plant protection agents’ agriculture is the main culprit, which doesn’t exclude certain pesticides being present in urban wastewater discharges.

As regards investigative monitoring, one should think about taking samples for laboratory analysis going upstream, step-by-step.

Of course, provided that the Water Quality Laboratory will be capable of analysing the relevant Priority substances and at sufficient low LODs.

Table 4-6 Likely sources of the selected Priority substances

| Priority substance | Likely sources |
| --- | --- |
| Chlorfenvinphos Chlorpyrifos Dieldrin Diuron Heptachlor epoxide A Terbutryn Tributyltin compounds (TBT) | (Plant protection products and biocides) - Agricultural activities - TBT: other sources |
| Fluoranthene Mercury Nickel Nonylphenols PFOS | - Urban wastewater discharges - Industrial activities |

# Conclusions

The Research monitoring programme has provided decent insight in ‘Priority and certain other pollutants’ in the Moldovan Dniester River Basin District. Actual field data are now available for virtually the full range of Priority substances in water. Not only at the nine sampling sites; the results allow for making reasonable assumptions at wider scales.

Sixty two (62) individual Priority substances mentioned in the Directive 2013/39/EU were analysed, out of which:

- 34 were not detected in any of the samples taken at the nine sites in April, July and October 2025;

- 16 substances were detected with concentrations less than 0.66 * AA-EQS of the Directive 2013/39/EU.

These 50 substances are considered to be of low priority; implying that there is no need for focussing on them, at least not in short/medium term.

Twelve (12) Priority substances are considered to be of short/medium term attention:

- 3 substances were found with concentrations exceeding the Maximum Allowable Concentration environmental quality standard (MAC-EQS) of the Directive 2013/39/EU. One concentration exceeding the MAC-EQS suffices to qualify a surface waterbody as ‘failing to achieve good chemical status’.

- 9 substances were detected with concentrations near to or higher than the Annual Average environmental quality standard of the Directive 2013/39/EU, at some of the sites. Three samples are though insufficient a number of data to draw conclusions about the chemical status.

Some suggestions for follow-up were made in Chapter 4 and in the Final Report of the Environmental Institute.

Generally, it is recommended to further discuss, assess and evaluate the results of this Research monitoring programme in dedicated expert groups.

Table 5-1 Detected Priority substances considered of short/medium term attention

|  | DNIESTER |  |  |  | RĂUT |  |  | ICHEL |  | BÎC |  |  | BOTNA |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Substance | Otaci | Nimereuca | Palanca |  | Bălţi | Ustia |  | Goian |  | Străşeni | Gura Bîcului |  | Chircăiești |
| Mercury |  |  |  |  |  |  |  |  |  |  | √ |  | √ |
| Chlorpyrifos |  |  |  |  | √ |  |  |  |  |  |  |  |  |
| Tributyltin compounds | √ | √ | √ |  |  |  |  |  |  |  |  |  |  |
| Nickel |  |  |  |  |  |  |  | √ |  | √ | √ |  | √ |
| Heptachlor epoxide A (*) |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
| PFOS |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
| Nonylphenols |  |  |  |  | √ |  |  |  |  |  | √ |  |  |
| Chlorfenvinphos |  |  |  |  |  |  |  |  |  |  | √ |  |  |
| Dieldrin |  |  |  |  |  |  |  |  |  |  | √ |  |  |
| Diuron |  |  |  |  |  |  |  |  |  |  | √ |  |  |
| Terbutryn |  |  |  |  |  |  |  |  |  |  | √ |  |  |
| Fluoranthene |  |  |  |  | √ |  |  |  |  |  |  |  |  |

(*) Concentrations were just above the LOD, hence results should be interpreted with care.

| √ | one or more concentrations >MAC-EQS |
| --- | --- |
| √ | one or more concentrations >AA-EQS |
| √ | one or more concentrations >0.66 * AA-EQS |

Annex 1	Environmental Quality Standards for Priority Substances and certain other pollutants

The table further below contains the Environmental Quality Standards (EQS) for Priority Substances and certain other Pollutants included in Annex II of the DIRECTIVE 2013/39/EU OF THE EUROPEAN PARLIAMENT AND OF THE COUNCIL of 12 August 2013 amending Directives 2000/60/EC and 2008/105/EC as regards priority substances in the field of water policy.

The water EQS laid down in this Annex are expressed as total concentrations in the whole water sample.

In the case of cadmium, lead, mercury and nickel (hereinafter “metals”), the water EQS refer to the dissolved concentration, i.e. the dissolved phase of a water sample obtained by filtration through a 0.45 μm filter or any equivalent pre-treatment, or, where specifically indicated, to the bioavailable concentration.

Member States may, when assessing the monitoring results against the relevant EQS, take into account: (a) natural background concentrations for metals and their compounds where such concentrations prevent compliance with the relevant EQS; (b) hardness, pH, dissolved organic carbon or other water quality parameters that affect the bioavailability of metals, the bioavailable concentrations being determined using appropriate bioavailability modelling.

ANNEX II

‘ANNEX I

ENVIRONMENTAL QUALITY STANDARDS FOR PRIORITY SUBSTANCES AND CERTAIN OTHER POLLUTANTS

PART A: ENVIRONMENTAL QUALITY STANDARDS (EQS)

AA: 	annual average.

MAC:	maximum allowable concentration.

Unit: 	[μg/l] for columns (4) to (7)

[μg/kg wet weight] for column (8)

| (1) | (2) | (3) | (4) | (5) | (6) | (7) | (8) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| No | Name of substance | CAS number (1) | AA-EQS (2) Inland surface waters (3) | AA-EQS (2) Other surface waters | MAC-EQS (4) Inland surface waters (3) | MAC-EQS (4) Other surface waters | EQS Biota(12) |
| (1) | Alachlor | 15972-60-8 | 0.3 | 0.3 | 0.7 | 0.7 |  |
| (2) | Anthracene | 120-12-7 | 0.1 | 0.1 | 0.1 | 0.1 |  |
| (3) | Atrazine | 1912-24-9 | 0.6 | 0.6 | 2.0 | 2.0 |  |
| (4) | Benzene | 71-43-2 | 10 | 8 | 50 | 50 |  |
| (5) | Brominated diphenylethers (5) | 32534-81-9 |  |  | 0.14 | 0.014 | 0.0085 |
| (6) | Cadmium and its compounds (depending on water hardness classes) (6) | 7440-43-9 | ≤ 0.08 (Class 1) 0.08 (Class 2) 0.09 (Class 3) 0.15 (Class 4) 0.25 (Class 5) | 0.2 | ≤ 0.45 (Class 1) 0.45 (Class 2) 0.6 (Class 3) 0.9 (Class 4) 1.5 (Class 5) | ≤ 0.45 (Class 1) 0.45 (Class 2) 0.6 (Class 3) 0.9 (Class 4) 1.5 (Class 5) |  |
| (6a) | Carbontetrachloride (7) | 56-23-5 | 12 | 12 | not applicable | not applicable |  |
| (7) | C10-13 Chloroalkanes (8) | 85535-84-8 | 0.4 | 0.4 | 1.4 | 1.4 |  |
| (8) | Chlorfenvinphos | 470-90-6 | 0.1 | 0.1 | 0.3 | 0.3 |  |
| (9) | Chlorpyrifos (Chlorpyrifos- ethyl) | 2921-88-2 | 0.03 | 0.03 | 0.1 | 0.1 |  |
| (9a) | Cyclodiene pesticides: Aldrin (7) Dieldrin (7) Endrin (7) Isodrin (7) | 309-00-2 60-57-1 72-20-8 465-73-6 | Σ = 0.01 | Σ = 0.005 | not applicable | not applicable |  |
| (9b) | DDT total (7), (9) | not applicable | 0.025 | 0.025 | not applicable | not applicable |  |
|  | para-para-DDT (7) | 50-29-3 | 0.01 | 0.01 | not applicable | not applicable |  |
| (10) | 1,2-Dichloroethane | 107-06-2 | 10 | 10 | not applicable | not applicable |  |
| (11) | Dichloromethane | 75-09-2 | 20 | 20 | not applicable | not applicable |  |
| (12) | Di(2- ethylhexyl)- phthalate (DEHP) | 117-81-7 | 1.3 | 1.3 | not applicable | not applicable |  |
| (13) | Diuron | 330-54-1 | 0.2 | 0.2 | 1.8 | 1.8 |  |
| (14) | Endosulfan | 115-29-7 | 0.005 | 0.0005 | 0.01 | 0.004 |  |
| (15) | Fluoranthene | 206-44-0 | 0.0063 | 0.0063 | 0.12 | 0.12 | 30 |
| (16) | Hexachlorobenzene | 118-74-1 |  |  | 0.05 | 0.05 | 10 |
| (17) | Hexachlorobutadiene | 87-68-3 |  |  | 0.6 | 0.6 | 55 |
| (18) | Hexachlorocyclohexane | 608-73-1 | 0.02 | 0.002 | 0.04 | 0.02 |  |
| (19) | Isoproturon | 34123-59-6 | 0.3 | 0.3 | 1.0 | 1.0 |  |
| (20) | Lead and its compounds | 7439-92-1 | 1.2 (13) | 1.3 | 14 | 14 |  |
| (21) | Mercury and its compounds | 7439-97-6 |  |  | 0.07 | 0.07 | 20 |
| (22) | Naphthalene | 91-20-3 | 2 | 2 | 130 | 130 |  |
| (23) | Nickel and its compounds | 7440-02-0 | 4 (13) | 8.6 | 34 | 34 |  |
| (24) | Nonylphenols (4-Nonylphenol) | 84852-15-3 | 0.3 | 0.3 | 2.0 | 2.0 |  |
| (25) | Octylphenols ((4-(1,1′,3,3′- tetramethylbutyl)-phenol)) | 140-66-9 | 0.1 | 0.01 | not applicable | not applicable |  |
| (26) | Pentachlorobenzene | 608-93-5 | 0.007 | 0.0007 | not applicable | not applicable |  |
| (27) | Pentachlorophenol | 87-86-5 | 0.4 | 0.4 | 1 | 1 |  |
| (28) | Polyaromatic hydrocarbons (PAH) (11) | not applicable | not applicable | not applicable | not applicable | not applicable |  |
|  | Benzo(a)pyrene | 50-32-8 | 1.7 × 10–4 | 1.7 × 10–4 | 0.27 | 0.027 | 5 |
|  | Benzo(b)fluoranthene | 205-99-2 | see footnote 11 | see footnote 11 | 0.017 | 0.017 | see footnote 11 |
|  | Benzo(k)fluoranthene | 207-08-9 | see footnote 11 | see footnote 11 | 0.017 | 0.017 | see footnote 11 |
|  | Benzo(g,h,i)perylene | 191-24-2 | see footnote 11 | see footnote 11 | 8.2 × 10–3 | 8.2 × 10–4 | see footnote 11 |
|  | Indeno(1,2,3- cd)pyrene | 193-39-5 | see footnote 11 | see footnote 11 | not applicable | not applicable | see footnote 11 |
| (29) | Simazine | 122-34-9 | 1 | 1 | 4 | 4 |  |
| (29a) | Tetrachloroethylene (7) | 127-18-4 | 10 | 10 | not applicable | not applicable |  |
| (29b) | Trichloroethylene (7) | 79-01-6 | 10 | 10 | not applicable | not applicable |  |
| (30) | Tributyltin compounds (Tributyltin- cation) | 36643-28-4 | 0.0002 | 0.0002 | 0.0015 | 0.0015 |  |
| (31) | Trichlorobenzenes | 12002-48-1 | 0.4 | 0.4 | not applicable | not applicable |  |
| (32) | Trichloromethane | 67-66-3 | 2.5 | 2.5 | not applicable | not applicable |  |
| (33) | Trifluralin | 1582-09-8 | 0.03 | 0.03 | not applicable | not applicable |  |
| (34) | Dicofol | 115-32-2 | 1.3 × 10–3 | 3.2 × 10–5 | not appli­ cable (10) | not appli­ cable (10) | 33 |
| (35) | Perfluorooctane sulfonic acid and its derivatives (PFOS) | 1763-23-1 | 6.5 × 10–4 | 1.3 × 10–4 | 36 | 7.2 | 9.1 |
| (36) | Quinoxyfen | 124495-18-7 | 0.15 | 0.015 | 2.7 | 0.54 |  |
| (37) | Dioxins and dioxin-like compounds | See footnote 10 in Annex X to Directive 2000/60/EC |  |  | not applicable | not applicable | Sum of PCDD+PCDF+ PCB-DL 0.0065 μg.kg–1 TEQ (14) |
| (38) | Aclonifen | 74070-46-5 | 0.12 | 0.012 | 0.12 | 0.012 |  |
| (39) | Bifenox | 42576-02-3 | 0.012 | 0.0012 | 0.04 | 0.004 |  |
| (40) | Cybutryne | 28159-98-0 | 0.0025 | 0.0025 | 0.016 | 0.016 |  |
| (41) | Cypermethrin | 52315-07-8 | 8 × 10–5 | 8 × 10–6 | 6 × 10–4 | 6 × 10–5 |  |
| (42) | Dichlorvos | 62-73-7 | 6 × 10–4 | 6 × 10–5 | 7 × 10–4 | 7 × 10–5 |  |
| (43) | Hexabromocyclododecane (HBCDD) | See footnote 12 in Annex X to Directive 2000/60/EC | 0.0016 | 0.0008 | 0.5 | 0.05 | 167 |
| (44) | Heptachlor and heptachlor epoxide | 76-44- 8/1024-57-3 | 2 × 10–7 | 1 × 10–8 | 3 × 10–4 | 3 × 10–5 | 6.7 × 10 –3 |
| (45) | Terbutryn | 886-50-0 | 0.065 | 0.0065 | 0.34 | 0.034 |  |

(1)  CAS: Chemical Abstracts Service.

(2)  This parameter is the EQS expressed as an annual average value (AA-EQS). Unless otherwise specified, it applies to the total concentration of all isomers.

(3)  Inland surface waters encompass rivers and lakes and related artificial or heavily modified water bodies.

(4)  This parameter is the EQS expressed as a maximum allowable concentration (MAC-EQS). Where the MAC-EQS are marked as “not applicable”, the AA-EQS values are considered protective against short-term pollution peaks in continuous discharges since they are significantly lower than the values derived on the basis of acute toxicity.

(5)  For the group of priority substances covered by brominated diphenylethers (No 5), the EQS refers to the sum of the concentrations of congener numbers 28, 47, 99, 100, 153 and 154.

(6)  For Cadmium and its compounds (No 6) the EQS values vary depending on the hardness of the water as specified in five class categories (Class 1: <40 mg CaCO3/l, Class 2: 40 to <50 mg CaCO3/l, Class 3: 50 to <100 mg CaCO3/l, Class 4: 100 to <200 mg CaCO3/l and Class 5: ≥ 200 mg CaCO3/l).

(7)  This substance is not a priority substance but one of the other pollutants for which the EQS are identical to those laid down in the legislation that applied prior to 13 January 2009.

(8)  No indicative parameter is provided for this group of substances. The indicative parameter(s) must be defined through the analytical method.

(9)  DDT total comprises the sum of the isomers 1,1,1-trichloro-2,2 bis (p-chlorophenyl) ethane (CAS number 50-29-3; EU number 200-024-3); 1,1,1-trichloro-2 (o-chlorophenyl)-2-(p-chlorophenyl) ethane (CAS number 789-02-6; EU Number 212-332-5); 1,1-dichloro-2,2 bis (p-chlorophenyl) ethylene (CAS number 72-55-9; EU Number 200-784-6); and 1,1-dichloro-2,2 bis (p-chlorophenyl) ethane (CAS number 72-54-8; EU Number 200-783-0).

(10) There is insufficient information available to set a MAC-EQS for these substances.

(11) For the group of priority substances of polyaromatic hydrocarbons (PAH) (No 28), the biota EQS and corresponding AA-EQS in water refer to the concentration of benzo(a)pyrene, on the toxicity of which they are based. Benzo(a)pyrene can be considered as a marker for the other PAH, hence only benzo(a)pyrene needs to be monitored for comparison with the biota EQS or the corresponding AA- EQS in water.

(12) Unless otherwise indicated, the biota EQS relate to fish. An alternative biota taxon, or another matrix, may be monitored instead, as long as the EQS applied provides an equivalent level of protection. For substances numbered 15 (Fluoranthene) and 28 (PAH), the biota EQS refers to crustaceans and molluscs. For the purpose of assessing chemical status, monitoring of Fluoranthene and PAH in fish is not appropriate. For substance number 37 (Dioxins and dioxin-like compounds), the biota EQS relates to fish, crustaceans and molluscs, in line with section 5.3 of the Annex to Commission Regulation (EU) No 1259/2011 of 2 December 2011 amending Regulation (EC) No 1881/2006 as regards maximum levels for dioxins, dioxin-like PCBs and non-dioxin-like PCBs in foodstuffs (OJ L 320, 3.12.2011, p. 18).

(13) These EQS refer to bioavailable concentrations of the substances.

(14) PCDD: polychlorinated dibenzo-p-dioxins; PCDF: polychlorinated dibenzofurans; PCB-DL: dioxin-like polychlorinated biphenyls; TEQ: toxic equivalents according to the World Health Organisation 2005 Toxic Equivalence Factors.’

Annex 2	In situ measured parameters

|  |  | DNIESTER |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Otaci |  |  | Nimereuca |  |  | Palanca |  |  |
| Parameter | Unit | Apr | Jul | Oct | Apr | Jul | Oct | Apr | Jul | Oct |
| Conductivity | uS/cm | 480 | 481 | 421 | 450 | 445 | 423 | 521 | 520 | 489 |
| pH | - | 4.7 | 7.9 | 7.2 | 8.5 | 8.7 | 7.5 | 7.7 | 7.5 | 7.2 |
| Hardness | mg CaCO3/L | 205 | 196 | 183 | 175 | 267 | 189 | 209 | 198 | 203 |
| Temperature | oC | 15.7 | 18.3 | 14.8 | 16.0 | 25.2 | 13.9 | 14.2 | 27.9 | 13.3 |
| O2 dissolved | mg O2/L | 3.7 | 10.2 | 7.2 | 4.2 | 13.6 | 7.7 | 6.5 | 6.8 | 6.5 |
| O2 saturation | % | 34 | 108 | 72 | 43 | 164 | 74 | 65 | 85 | 63 |

|  |  | RĂUT |  |  |  |  |  |  | ICHEL |  |  |  | BÎC |  |  |  |  |  |  | BOTNA |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Bălţi |  |  | Ustia |  |  |  | Goian |  |  |  | Străşeni |  |  | Gura Bîcului |  |  |  | Chircăiești |  |  |
| Parameter | Unit | Apr | Jul | Oct | Apr | Jul | Oct |  | Apr | Jul | Oct |  | Apr | Jul | Oct | Apr | Jul | Oct |  | Apr | Jul | Oct |
| Conductivity | µS/cm | 1734 | 1416 | 1490 | 1857 | 1816 | 1723 |  | 1739 | 1602 | 1646 |  | 1790 | 1988 | 2070 | 1568 | 1165 | 1263 |  | 3260 | 6110 | 3360 |
| pH | - | 7.5 | 7.7 | 7.9 | 8.2 | 8.2 | 8.2 |  | 7.5 | 7.1 | 7.5 |  | 7.5 | 7.8 | 8.1 | 7.5 | 7.3 | 7.6 |  | 8.5 | 9.1 | 9.2 |
| Hardness | mg CaCO3/L | 287 | 280 | 245 | 261 | 270 | 266 |  | 313 | >350 | 290 |  | 128 | 330 | 290 | 285 | >350 | 316 |  | 321 | >350 | 316 |
| Temperature | oC | 17.5 | 23.9 | 17.0 | 16.1 | 25.5 | 9.9 |  | 13.0 | 18.6 | 10.3 |  | 10.9 | 23.0 | 12.2 | 19.1 | 28.0 | 15.5 |  | 20.1 | 28.3 | 13.3 |
| O2 dissolved | mg O2/L | 1.7 | 5.0 | 7.5 | 2.4 | 4.9 | 8.5 |  | 0.7 | 3.0 | 6.0 |  | 3.0 | 1.5 | 6.7 | 0.4 | 0.8 | 3.3 |  | 4.6 | 15.0 | 17.0 |
| O2 saturation | % | 17 | 62 | 77 | 24 | 62 | 75 |  | 7 | 32 | 56 |  | 30 | 19 | 63 | 4 | 10 | 31 |  | 51 | 190 | 164 |

Notes:

- In April 2025, the sample at Otaci was not taken from the bridge but from a small inlet at the right bank a few hundred meters downstream.

Annex 3	Limit of detection, limit of quantification, AA-EQS, MAC-EQS

| Group | Compound | Unit | LOD | LOQ | AA- EQS | MAC- EQS | Remarks |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HM | Cadmium and its compounds | µg/l | 0.03 | 0.10 | 0.25 | 1.5 | Class 5: ≥ 200 mg CaCO3/l |
|  | Lead and its compounds | µg/l | 0.03 | 0.1 | 1.2 | 14 |  |
|  | Mercury and its compounds | µg/l | 0.0061 | 0.02 | - | 0.07 |  |
|  | Nickel and its compounds | µg/l | 0.12 | 0.4 | 4 | 34 |  |
| OCP | 1,2,3,4,5,6-Hexachlorocyclohexane (HCH total) | ng/l | 0.004 | 0.01 | 20 | 40 |  |
|  | Aldrin | ng/l | 0.04 | 0.1 | 10 | - | ∑ drins = 10 |
|  | Chlorfenvinphos | ng/l | 0.04 | 0.1 | 100 | 300 |  |
|  | Chlorpyrifos (Chlorpyrifos-ethyl) | ng/l | 0.04 | 0.1 | 30 | 100 |  |
|  | DDT total | ng/l | 0.04 | 0.1 | 25 | - |  |
|  | Dichlorvos | ng/l | 0.04 | 0.1 | 0.6 | 0.7 |  |
|  | Dicofol | ng/l | 0.04 | 0.1 | 1.3 | - |  |
|  | Dieldrin | ng/l | 0.04 | 0.1 | 10 | - | ∑ drins = 10 |
|  | Endosulfan I | ng/l | 0.04 | 0.1 | 5 | 10 |  |
|  | Endosulfan II | ng/l | 0.04 | 0.1 | 5 | 10 |  |
|  | Endrin | ng/l | 0.04 | 0.1 | 10 | - | ∑ drins = 10 |
|  | Heptachlor | ng/l | 0.04 | 0.1 | 0.0002 | 0.3 |  |
|  | Heptachlor epoxide A | ng/l | 0.004 | 0.01 | 0.0002 | 0.3 |  |
|  | Heptachlor epoxide B | ng/l | 0.004 | 0.01 | 0.0002 | 0.3 |  |
|  | Hexachlorobenzene | ng/l | 0.04 | 0.1 | - | 50 |  |
|  | Isodrin | ng/l | 0.04 | 0.1 | 10 | - | ∑ drins = 10 |
|  | o,p'-DDT | ng/l | 0.04 | 0.1 | - | - |  |
|  | p,p'-DDD | ng/l | 0.04 | 0.1 | - | - |  |
|  | p,p'-DDE | ng/l | 0.04 | 0.1 | - | - |  |
|  | p,p'-DDT | ng/l | 0.04 | 0.1 | 10 | - |  |
|  | Pentachlorobenzene | ng/l | 0.04 | 0.1 | 7 | - |  |
|  | Pentachlorophenol | ng/l | 20 | 50 | 400 | 1000 |  |
| PEST | Aclonifen | ng/l | 0.61 | 1.84 | 120 | 120 |  |
|  | Alachlor | ng/l | 0.11 | 0.34 | 300 | 700 |  |
|  | Atrazine | ng/l | 15.67 | 47.00 | 600 | 2000 |  |
|  | Bifenox | ng/l | 0.65 | 1.95 | 12 | 40 |  |
|  | Cybutryne | ng/l | 0.94 | 2.82 | 2.5 | 16 |  |
|  | Cypermethrin | ng/l | 0.16 | 0.49 | 0.08 | 0.6 |  |
|  | Diuron | ng/l | 0.37 | 1.12 | 200 | 1800 |  |
|  | Isoproturon | ng/l | 0.06 | 0.17 | 300 | 1000 |  |
|  | Quinoxyfen | ng/l | 0.59 | 1.76 | 150 | 2700 |  |
|  | Simazine | ng/l | 6.65 | 19.94 | 1000 | 4000 |  |
|  | Terbutryn | ng/l | 1.33 | 3.98 | 65 | 340 |  |
|  | Trifluralin | ng/l | 0.04 | 0.1 | 30 | - |  |
| PAH | Anthracene | ng/l | 0.03 | 0.1 | 100 | 100 |  |
|  | Benzo(a)pyrene | ng/l | 0.03 | 0.1 | 0.17 | 270 |  |
|  | Benzo(b)fluoranthene | ng/l | 0.03 | 0.1 | - | 17 |  |
|  | Benzo[g,h,i]perylene | ng/l | 0.03 | 0.1 | - | 8.2 |  |
|  | Benzo[k]fluoranthene | ng/l | 0.03 | 0.1 | - | 17 |  |
|  | Fluoranthene | ng/l | 0.03 | 0.1 | 6.3 | 120 |  |
|  | Indeno(1,2,3-cd)pyrene | ng/l | 0.03 | 0.1 | - | - |  |
| VOC | 1,2-dichloroethane | µg/l | 0.32 | 1 | 10 | - |  |
|  | Aniline | µg/l | 0.05 | 0.1 | -- | - |  |
|  | Benzene | µg/l | 0.04 | 0.1 | 10 | 50 |  |
|  | Carbon-tetrachloride | µg/l | 0.035 | 0.1 | 12 | - |  |
|  | Dichloromethane | µg/l | 0.33 | 1 | 20 | -- |  |
|  | Hexachlorobutadiene | µg/l | 0.032 | 0.1 |  | 0.6 |  |
|  | Naphtalene | µg/l | 0.004 | 0.01 | 2 | 130 |  |
|  | Styrene | µg/l | 0.04 | 0.1 | - | - |  |
|  | Tetrachloro- ethylene | µg/l | 0.1 | 1 | 10 | - |  |
|  | Toluene | µg/l | 0.04 | 0.1 | - | - |  |
|  | Trichloro- ethylene | µg/l | 0.15 | 1 | 10 | - |  |
|  | Trichlorobenzenes | µg/l | 0.03 | 0.1 | 0.4 | - |  |
|  | Trichloromethane (chloroform) | µg/l | 0.03 | 0.1 | 2.5 | - |  |
| IND | 4-Nonylphenol | µg/l | 0.02 | 0.05 | 0.3 | 2 |  |
|  | 4-Octylphenol | µg/l | 0.02 | 0.05 | 0.1 | - |  |
|  | Brominated diphenylethers | ng/l | 0.03 | 0.1 | - | 140 |  |
|  | Chloroalkanes, C10-13 | ng/l | 5 | 20 | 400 | 1400 |  |
|  | Di(2-ethylhexyl)phthalate (DEHP) | µg/l | 0.04 | 0.1 | 1.3 | - |  |
|  | Hexabromocyclododecane (HBCDD) | ng/l | 0.03 | 0.1 | 1.6 | 500 |  |
|  | Nonylphenols | µg/l | 0.02 | 0.05 | 0.3 | 2 |  |
|  | Octylphenols ((4-(1,1’,3, 3’-tetramethyl-butyl)-phenol)) | µg/l | 0.02 | 0.05 | 0.1 | - |  |
| TBT | Tributyltin compounds | ng/l | 0.04 | 0.1 | 0.2 | 1 |  |
| PFOS | Linear_ Perfluorooctanesulfonic acid (PFOS) | ng/l | 0.02 | 0.06 | - | - |  |
|  | Total_ Perfluorooctanesulfonic acid (PFOS) | ng/l | 0.02 | 0.06 | 0.65 | - |  |

Annex 4	Substances with concentrations <LOD in all samples

LOQ ≥ 30% of EQS (LOD ≥ 10% of EQS

limit of detection >AA-EQS

| Group | Compound | Unit | LOD | LOQ | AA EQS | MAC EQS | Remarks |
| --- | --- | --- | --- | --- | --- | --- | --- |
| OCP | Aldrin | ng/l | 0.04 | 0.1 | 10 |  | ∑ drins = 10 |
|  | Dichlorvos | ng/l | 0.04 | 0.1 | 0.6 | 0 7 |  |
|  | Dicofol | ng/l | 0.04 | 0.1 | 1.3 |  |  |
|  | Endosulfan I | ng/l | 0.04 | 0.1 | 5 | 10 |  |
|  | Endosulfan II | ng/l | 0.04 | 0.1 | 5 | 10 |  |
|  | Endrin | ng/l | 0.04 | 0.1 | 10 |  | ∑ drins = 10 |
|  | Heptachlor | ng/l | 0.04 | 0.1 | 0.0002 | 0.3 |  |
|  | Heptachlor epoxide B | ng/l | 0.004 | 0.01 | 0.0002 | 0.3 |  |
|  | Isodrin | ng/l | 0.04 | 0.1 | 10 |  | ∑ drins = 10 |
|  | o,p'-DDT | ng/l | 0.04 | 0.1 |  |  |  |
|  | p,p'-DDT | ng/l | 0.04 | 0.1 | 10 |  |  |
|  | Pentachlorobenzene | ng/l | 0.04 | 0.1 | 7 |  |  |
|  | Pentachlorophenol | ng/l | 20 | 50 | 400 | 1000 |  |
| PEST | Aclonifen | ng/l | 0.61 | 1.84 | 120 | 120 |  |
|  | Alachlor | ng/l | 0.11 | 0.34 | 300 | 700 |  |
|  | Bifenox | ng/l | 0.65 | 1.95 | 12 | 40 |  |
|  | Cybutryne | ng/l | 0.94 | 2.82 | 2.5 | 16 |  |
|  | Cypermethrin | ng/l | 0.16 | 0.49 | 0.08 | 0 6 |  |
|  | Quinoxyfen | ng/l | 0.59 | 1.76 | 150 | 2700 |  |
|  | Trifluralin | ng/l | 0.04 | 0.1 | 30 |  |  |
| PAH | Benzo(a)pyrene | ng/l | 0.03 | 0.1 | 0.17 | 270 |  |
|  | Benzo[k]fluoranthene | ng/l | 0.03 | 0.1 |  | 17 |  |
| VOC | 1,2-dichloroethane | µg/l | 0.32 | 1 | 10 |  |  |
|  | Aniline | µg/l | 0.05 | 0.1 |  |  |  |
|  | Benzene | µg/l | 0.04 | 0.1 | 10 | 50 |  |
|  | Carbon-tetrachloride | µg/l | 0.035 | 0.1 | 12 |  |  |
|  | Dichloromethane | µg/l | 0.33 | 1 | 20 |  |  |
|  | Hexachlorobutadiene | µg/l | 0.032 | 0.1 |  | 0 6 |  |
|  | Styrene | µg/l | 0.04 | 0.1 |  |  |  |
|  | Tetrachloro- ethylene | µg/l | 0.1 | 1 | 10 |  |  |
|  | Trichloro- ethylene | µg/l | 0.15 | 1 | 10 |  |  |
|  | Trichlorobenzenes | µg/l | 0.03 | 0.1 | 40 |  |  |
| IND | 4-Nonylphenol | µg/l | 0.02 | 0.05 | 0.3 | 2 |  |
|  | 4-Octylphenol | µg/l | 0.02 | 0.05 | 0.1 |  |  |
|  | Brominated diphenylethers | ng/l | 0.03 | 0.1 |  | 140 |  |
|  | Hexabromocyclododecane (HBCDD) | ng/l | 0.03 | 0.1 | 1.6 | 500 |  |
|  | Octylphenols ((4-(1,1’,3, 3’-tetramethyl-butyl)-phenol)) | µg/l | 0.02 | 0.05 | 0.1 |  |  |

Annex 5	Overviews with concentrations ≥LOD in one or more samples

Notice that the tables below were generated from the same template. For example, cadmium was detected only in the July Răut-Bălţi sample. Nevertheless, cadmium is mentioned in all tables.

|  |  |  | DNIESTER |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Otaci |  |  | Nimereuca |  |  | Palanca |  |  |  |  |  |  |
| Group | Substance name | Unit | Apr | Jul | Oct | Apr | Jul | Oct | Apr | Jul | Oct |  | LOD | AA-EQS | MAC-EQS |
| HM | Cadmium and its compounds | µg/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.03 | 0.25 | 1.5 |
|  | Lead and its compounds | µg/l | 0.243 | 0.14 | 0.20 | <LOD | 0.088 | <LOD | <LOD | 0.032 | <LOD |  | 0.03 | 1.2 | 14 |
|  | Mercury and its compounds | µg/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.0061 |  | 0.07 |
|  | Nickel and its compounds | µg/l | <LOD | 0.9 | 1.78 | <LOD | 0.8 | 0.96 | <LOD | 1.4 | 0.99 |  | 0.12 | 4 | 34 |
| OCP | 1,2,3,4,5,6-Hexachlorocyclohexane (HCH total) | ng/l | 1.76 | <LOD | 0.21 | 1.88 | <LOD | 0.19 | 3.92 | <LOD | 1.33 |  | 0.004 | 20 | 40 |
|  | Chlorfenvinphos | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.04 | 100 | 300 |
|  | Chlorpyrifos (Chlorpyrifos-ethyl) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.04 | 30 | 100 |
|  | DDT total | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.04 | 25 |  |
|  | Dieldrin | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.04 | 10 |  |
|  | Heptachlor epoxide A | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.004 | 0.0002 | 0.3 |
|  | Hexachlorobenzene | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.04 |  | 50 |
|  | p,p'-DDD | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.04 |  |  |
|  | p,p'-DDE | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.04 |  |  |
| PEST | Atrazine | ng/l | 47.20 | <LOD | <LOD | 47.48 | <LOD | <LOD | 32.11 | <LOD | <LOD |  | 15.67 | 600 | 2000 |
|  | Diuron | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 3.10 | 2.8 | 2.06 |  | 0.37 | 200 | 1800 |
|  | Isoproturon | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.06 | 300 | 1000 |
|  | Simazine | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 6.65 | 1000 | 4000 |
|  | Terbutryn | ng/l | 4.37 | <LOD | 14.7 | <LOD | <LOD | 15.4 | <LOD | 4.4 | 4.14 |  | 1.33 | 65 | 340 |
| PAH | Anthracene | ng/l | 16.22 | 1.22 | 0.56 | 12.55 | 0.92 | 0.51 | <LOD | <LOD | <LOD |  | 0.03 | 100 | 100 |
|  | Benzo(b)fluoranthene | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.03 |  | 17 |
|  | Benzo[g,h,i]perylene | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.03 |  | 8.2 |
|  | Fluoranthene | ng/l | 2.55 | 0.91 | <LOD | 2.1 | 0.85 | <LOD | <LOD | <LOD | <LOD |  | 0.03 | 6.3 | 120 |
|  | Indeno(1,2,3-cd)pyrene | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.03 |  |  |
| VOC | Naphtalene | µg/l | 0.44 | 0.006 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.004 | 2 | 130 |
|  | Toluene | µg/l | 0.22 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 0.044 | <LOD |  | 0.04 |  |  |
|  | Trichloromethane (chloroform) | µg/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.03 | 2.5 |  |
| IND | Chloroalkanes, C10-13 | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 5 | 400 | 1400 |
|  | Di(2-ethylhexyl)phthalate (DEHP) | µg/l | <LOD | <LOD | 0.22 | <LOD | <LOD | 0.19 | 0.12 | <LOD | 0.13 |  | 0.04 | 1.3 |  |
|  | Nonylphenols | µg/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 0.033 | <LOD | <LOD |  | 0.02 | 0.3 | 2 |
| TBT | Tributyltin compounds | ng/l | 43.30 | 1.3 | 7.86 | 26.2 | 1.02 | 6.81 | 1.22 | 0.044 | 1.01 |  | 0.04 | 0.2 | 1.5 |
| PFOS | Linear_ Perfluorooctanesulfonic acid (PFOS) | ng/l | 0.028 | 0.227 | <LOD | 0.027 | 0.177 | <LOD | 0.027 | 0.146 | <LOD |  | 0.02 | 0.65 | 36000 |
|  | Total_ Perfluorooctanesulfonic acid (PFOS) | ng/l | 0.14 | 0.297 | <LOD | 0.22 | 0.249 | <LOD | 0.10 | 0.204 | <LOD |  | 0.02 | 0.65 | 36000 |

Notes:

- In April 2025, the sample at Otaci was not taken from the bridge but from a small inlet at the right bank a few hundred meters downstream.

|  |  |  |  | RĂUT |  |  |  |  |  |  |  | ICHEL |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Bălţi |  |  |  | Ustia |  |  |  | Goian |  |  |  |  |  |  |
| Group | Substance name | Unit |  | Apr | Jul | Oct |  | Apr | Jul | Oct |  | Apr | Jul | Oct |  | LOD | AA-EQS | MAC-EQS |
| HM | Cadmium and its compounds | µg/l |  | <LOD | 0.041 | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.03 | 0.25 | 1.5 |
|  | Lead and its compounds | µg/l |  | 0.107 | 0.22 | 0.14 |  | <LOD | 0.075 | <LOD |  | 0.278 | 0.10 | 0.06 |  | 0.03 | 1.2 | 14 |
|  | Mercury and its compounds | µg/l |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.011 | <LOD | <LOD |  | 0.0061 | - | 0.07 |
|  | Nickel and its compounds | µg/l |  | <LOD | 1.9 | 3.27 |  | 0.152 | 2.6 | 2.56 |  | <LOD | 2.0 | 4.99 |  | 0.12 | 4 | 34 |
| OCP | 1,2,3,4,5,6-Hexachlorocyclohexane (HCH total) | ng/l |  | 3.22 | 4.22 | 1.22 |  | 2.49 | <LOD | 0.89 |  | 9.82 | <LOD | 1.56 |  | 0.004 | 20 | 40 |
|  | Chlorfenvinphos | ng/l |  | <LOD | 5.22 | 1.22 |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 | 100 | 300 |
|  | Chlorpyrifos (Chlorpyrifos-ethyl) | ng/l |  | 36.73 | 135 | 25.2 |  | <LOD | <LOD | <LOD |  | 2.19 | <LOD | 1.45 |  | 0.04 | 30 | 100 |
|  | DDT total | ng/l |  | <LOD | 0.21 | <LOD |  | <LOD | <LOD | <LOD |  | 0.1 | <LOD | <LOD |  | 0.04 | 25 | - |
|  | Dieldrin | ng/l |  | <LOD | 0.88 | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 | 10 | - |
|  | Heptachlor epoxide A | ng/l |  | <LOD | 0.005 | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.004 | 0.0002 | 0.3 |
|  | Hexachlorobenzene | ng/l |  | <LOD | 0.08 | <LOD |  | <LOD | 0.05 | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 | - | 50 |
|  | p,p'-DDD | ng/l |  | <LOD | 0.06 | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 | - | - |
|  | p,p'-DDE | ng/l |  | <LOD | 0.21 | <LOD |  | <LOD | <LOD | <LOD |  | 0.1 | <LOD | <LOD |  | 0.04 | - | - |
| PEST | Atrazine | ng/l |  | 122.28 | <LOD | <LOD |  | 82.58 | 71 | <LOD |  | 97.12 | <LOD | <LOD |  | 15.67 | 600 | 2000 |
|  | Diuron | ng/l |  | 45.28 | 24 | 29.9 |  | 1.41 | 0.88 | 0.55 |  | 41.64 | 29 | 29.0 |  | 0.37 | 200 | 1800 |
|  | Isoproturon | ng/l |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.06 | 300 | 1000 |
|  | Simazine | ng/l |  | <LOD | <LOD | <LOD |  | 23.92 | <LOD | <LOD |  | 78.06 | <LOD | <LOD |  | 6.65 | 1000 | 4000 |
|  | Terbutryn | ng/l |  | 27.73 | <LOD | <LOD |  | 3.22 | <LOD | 2.66 |  | 17.49 | <LOD | <LOD |  | 1.33 | 65 | 340 |
| PAH | Anthracene | ng/l |  | 18.34 | 10.2 | 2.38 |  | <LOD | 2.12 | <LOD |  | 10.22 | 2.33 | 0.92 |  | 0.03 | 100 | 100 |
|  | Benzo(b)fluoranthene | ng/l |  | 1.22 | 1.32 | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.03 | - | 17 |
|  | Benzo[g,h,i]perylene | ng/l |  | 1.56 | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.03 | - | 8.2 |
|  | Fluoranthene | ng/l |  | 4.55 | 5.23 | 0.89 |  | <LOD | <LOD | <LOD |  | 2.44 | 1.32 | 0.51 |  | 0.03 | 6.3 | 120 |
|  | Indeno(1,2,3-cd)pyrene | ng/l |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.03 | - | - |
| VOC | Naphtalene | µg/l |  | <LOD | 0.008 | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.004 | 2 | 130 |
|  | Toluene | µg/l |  | 0.83 | <LOD | 0.27 |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 | - | - |
|  | Trichloromethane (chloroform) | µg/l |  | 0.88 | 1.09 | 0.42 |  | <LOD | <LOD | 0.13 |  | <LOD | 0.05 | <LOD |  | 0.03 | 2.5 | - |
| IND | Chloroalkanes, C10-13 | ng/l |  | 122 | 107 | 49.4 |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 5 | 400 | 1400 |
|  | Di(2-ethylhexyl)phthalate (DEHP) | µg/l |  | 0.22 | 0.75 | 0.68 |  | <LOD | <LOD | <LOD |  | <LOD | 0.22 | <LOD |  | 0.04 | 1.3 | - |
|  | Nonylphenols | µg/l |  | 0.290 | 0.86 | 0.21 |  | <LOD | <LOD | 0.09 |  | 0.160 | <LOD | <LOD |  | 0.02 | 0.3 | 2 |
| 1TBT | Tributyltin compounds | ng/l |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 | 0.2 | 1.5 |
| PFOS | Linear_ Perfluorooctanesulfonic acid (PFOS) | ng/l |  | 0.055 | 0.539 | <LOD |  | 0.051 | 0.055 | <LOD |  | 0.024 | 0.031 | <LOD |  | 0.02 | 0.65 | - |
|  | Total_ Perfluorooctanesulfonic acid (PFOS) | ng/l |  | 0.11 | 0.921 | <LOD |  | 0.18 | 0.034 | <LOD |  | 0.19 | 0.040 | <LOD |  | 0.02 | 0.65 | - |

|  |  |  | BIC |  |  |  |  |  |  | BOTNA |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Străşeni |  |  | Gura Bîcului |  |  |  | Chircăiești |  |  |  |  |  |  |  |
| Group | Substance name | Unit | Apr | Jul | Oct | Apr | Jul | Oct |  | Apr | Jul | Oct | . | LOD |  | AA-EQS | MAC-EQS |
| HM | Cadmium and its compounds | µg/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.03 |  | 0.25 | 1.5 |
|  | Lead and its compounds | µg/l | <LOD | 0.041 | 0.27 | 0.082 | 0.14 | 0.33 |  | 0.192 | 0.40 | 0.12 |  | 0.03 |  | 1.2 | 14 |
|  | Mercury and its compounds | µg/l | <LOD | <LOD | <LOD | 0.150 | <LOD | <LOD |  | 1.195 | <LOD | <LOD |  | 0.0061 |  | - | 0.07 |
|  | Nickel and its compounds | µg/l | <LOD | 2.9 | 5.35 | 0.213 | 2.5 | 2.23 |  | 1.327 | 6.7 | 3.95 |  | 0.12 |  | 4 | 34 |
| OCP | 1,2,3,4,5,6-Hexachlorocyclohexane (HCH total) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | 0.42 |  | 3.36 | <LOD | 2.02 |  | 0.004 |  | 20 | 40 |
|  | Chlorfenvinphos | ng/l | <LOD | <LOD | <LOD | <LOD | 16.9 | <LOD |  | <LOD | 95 | 1.05 |  | 0.04 |  | 100 | 300 |
|  | Chlorpyrifos (Chlorpyrifos-ethyl) | ng/l | <LOD | <LOD | <LOD | 16.83 | 12 | 9.88 |  | <LOD | <LOD | <LOD |  | 0.04 |  | 30 | 100 |
|  | DDT total | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | 2.976403 |  | 1.17 | 0.44 | 2.02 |  | 0.04 |  | 25 | - |
|  | Dieldrin | ng/l | <LOD | <LOD | <LOD | 7.65 | 0.56 | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 |  | 10 | - |
|  | Heptachlor epoxide A | ng/l | <LOD | <LOD | <LOD | <LOD | 0.006 | <LOD |  | <LOD | <LOD | <LOD |  | 0.004 |  | 0.0002 | 0.3 |
|  | Hexachlorobenzene | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 |  | - | 50 |
|  | p,p'-DDD | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | 0.46 |  | 0.88 | 0.32 | 0.52 |  | 0.04 |  | - | - |
|  | p,p'-DDE | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | 0.33 |  | 0.29 | 0.12 | 0.29 |  | 0.04 |  | - | - |
| PEST | Atrazine | ng/l | 57.40 | <LOD | <LOD | 101.22 | <LOD | <LOD |  | 77.86 | <LOD | 22.1 |  | 15.67 |  | 600 | 2000 |
|  | Diuron | ng/l | 0.85 | 9.0 | 41.1 | 192.02 | 113 | 71.9 |  | 6.69 | 4.5 | 4.44 |  | 0.37 |  | 200 | 1800 |
|  | Isoproturon | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | 2.90 |  | <LOD | <LOD | <LOD |  | 0.06 |  | 300 | 1000 |
|  | Simazine | ng/l | 22.02 | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 6.65 |  | 1000 | 4000 |
|  | Terbutryn | ng/l | 2.88 | <LOD | 1.78 | 48.49 | <LOD | <LOD |  | 19.89 | 28.3 | 12.1 |  | 1.33 |  | 65 | 340 |
| PAH | Anthracene | ng/l | <LOD | <LOD | <LOD | 5.89 | 4.65 | 3.11 |  | <LOD | <LOD | <LOD |  | 0.03 |  | 100 | 100 |
|  | Benzo(b)fluoranthene | ng/l | <LOD | <LOD | <LOD | <LOD | 0.99 | <LOD |  | <LOD | <LOD | <LOD |  | 0.03 |  | - | 17 |
|  | Benzo[g,h,i]perylene | ng/l | <LOD | <LOD | <LOD | <LOD | 1.22 | <LOD |  | <LOD | <LOD | <LOD |  | 0.03 |  | - | 8.2 |
|  | Fluoranthene | ng/l | <LOD | <LOD | <LOD | 3.56 | 2.66 | 0.42 |  | <LOD | <LOD | <LOD |  | 0.03 |  | 6.3 | 120 |
|  | Indeno(1,2,3-cd)pyrene | ng/l | <LOD | <LOD | <LOD | <LOD | 0.99 | <LOD |  | <LOD | <LOD | <LOD |  | 0.03 |  | - | - |
| VOC | Naphtalene | µg/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.004 |  | 2 | 130 |
|  | Toluene | µg/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 |  | - | - |
|  | Trichloromethane (chloroform) | µg/l | <LOD | <LOD | <LOD | 0.4 | 0.3 | 0.21 |  | <LOD | <LOD | <LOD |  | 0.03 |  | 2.5 | - |
| IND | Chloroalkanes, C10-13 | ng/l | <LOD | <LOD | <LOD | 65.22 | 94 | 11.1 |  | <LOD | <LOD | <LOD |  | 5 |  | 400 | 1400 |
|  | Di(2-ethylhexyl)phthalate (DEHP) | µg/l | <LOD | 0.15 | <LOD | <LOD | 0.34 | 0.33 |  | <LOD | 0.19 | <LOD |  | 0.04 |  | 1.3 | - |
|  | Nonylphenols | µg/l | <LOD | <LOD | <LOD | 0.130 | 0.23 | 0.08 |  | <LOD | <LOD | <LOD |  | 0.02 |  | 0.3 | 2 |
| TBT | Tributyltin compounds | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 |  | 0.2 | 1.5 |
| PFOS | Linear_ Perfluorooctanesulfonic acid (PFOS) | ng/l | 0.066 | 0.224 | <LOD | 0.19 | 1.770 | <LOD |  | 0.044 | 0.057 | <LOD |  | 0.02 |  | 0.65 | - |
|  | Total_ Perfluorooctanesulfonic acid (PFOS) | ng/l | 0.33 | 0.263 | <LOD | 0.44 | 2.418 | <LOD |  | 0.11 | 0.041 | <LOD |  | 0.02 |  | 0.65 | - |

Annex 6	Overview of the results of analysis for PFAS

Not detected in any of the samples

| Substance name |  | LOD ng/l |
| --- | --- | --- |
| 1H,1H,2H,2H-Perfluorodecanesulfonic acid (8:2FTS) |  | 0.22 |
| 1H,1H,2H,2H-Perfluorohexanesulfonic acid (4:2FTS) |  | 0.15 |
| 6:2 Fluorotelomer phosphate diester (6:2diPAP) |  | 0.05 |
| Bis(heptadecafluorooctyl)phosphinic acid (8:8PFPi) |  | 0.08 |
| Bis(tridecafluorohexyl)phosphinic acid (6-6PFPi) |  | 0.02 |
| Capstone B (DPOSA) |  | 0.10 |
| N-Ethyl-n-perfluorooctanesulfoneamide (N-EtPFOSA) |  | 0.13 |
| N-Ethyl-n-perfluorooctanesulfonylglycine (N-EtFOSAA) |  | 0.06 |
| N-Methyl-n-perfluoroctanesulfonamide (N-MePFOSA) |  | 0.02 |
| N-Methyl-n-perfluorooctanesulfonylglycine (N-MeFOSAA) |  | 0.06 |
| Perfluoro(2-((6-chlorohexyl)oxy)ethanesulfonic acid) (F53B-Major) |  | 0.02 |
| Perfluoro(2-((6-chlorohexyl)oxy)ethanesulfonic acid) (F53B-Minor) |  | 0.19 |
| Perfluoro-1-octanesulfonamidoacetic acid (FOSAA) |  | 0.05 |
| Perfluorodecanesulfonic acid (PFDS) |  | 0.06 |
| Perfluorohexadecanoic acid (PFHxDA) |  | 0.22 |
| Perfluorohexylperfluorooctyl phosphinic acid (6:8PFPi) |  | 0.15 |
| Perfluorotetradecanoic acid (PFTeDA) |  | 0.08 |
| Perfluorotridecanoic acid (PFTrDA) |  | 0.12 |

Detected in one or more samples

|  |  | DNIESTER |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Otaci |  |  | Nimereuca |  |  | Palanca |  |  |  |  |
| Substance name | Unit | Apr | Jul | Oct | Apr | Jul | Oct | Apr | Jul | Oct |  | LOD ng/l |
| 1H,1H,2H,2H-Perfluorooctanesulfonic acid (6:2FTS) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.17 |
| 2,3,3,3-Tetrafluoro-2(heptafluoropropoxy)propionic acid (GenX) | ng/l | 0.47 | 0.49 | 0.02 | 0.53 | 0.49 | 0.02 | 0.18 | 0.38 | <LOD |  | 0.02 |
| Capstone A (CDPOS) | ng/l | 0.35 | <LOD | 0.05 | 0.52 | <LOD | 0.05 | <LOD | <LOD | <LOD |  | 0.04 |
| Dodecafluoro-3H-4,8-dioxanonate (ADONA) | ng/l | <LOD | 0.122 | <LOD | <LOD | 0.115 | <LOD | <LOD | 0.172 | <LOD |  | 0.10 |
| Perfluorobutanesulfonic acid (PFBS) | ng/l | 1.2 | 0.311 | 0.12 | 4.4 | 0.499 | 0.11 | 0.54 | 0.321 | <LOD |  | 0.08 |
| Perfluorobutanoic acid (PFBA) | ng/l | <LOD | 0.125 | <LOD | <LOD | 0.177 | <LOD | <LOD | 0.116 | <LOD |  | 0.11 |
| Perfluorodecanoic acid (PFDA) | ng/l | <LOD | 0.090 | <LOD | <LOD | 0.092 | <LOD | <LOD | 0.105 | <LOD |  | 0.08 |
| Perfluorododecanoic acid (PFDoA) | ng/l | 0.012 | 0.069 | <LOD | 0.011 | 0.056 | <LOD | <LOD | 0.055 | <LOD |  | 0.01 |
| Perfluoroheptanesulfonic acid (PFHpS) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.14 |
| Perfluoroheptanoic acid (PFHpA) | ng/l | <LOD | 0.112 | <LOD | 1.4 | 0.121 | <LOD | 0.31 | 0.122 | <LOD |  | 0.10 |
| Perfluorohexanesulfonic acid (PFHxS) | ng/l | <LOD | 0.045 | <LOD | 1.2 | 0.041 | <LOD | 0.24 | 0.051 | <LOD |  | 0.04 |
| Perfluorohexanoic acid (PFHxA) | ng/l | 0.19 | 0.106 | <LOD | 0.25 | 0.128 | 0.04 | 0.031 | 0.034 | <LOD |  | 0.03 |
| Perfluorononanesulfonic acid (PFNS) | ng/l | 0.022 | <LOD | <LOD | 0.025 | <LOD | <LOD | 0.022 | <LOD | <LOD |  | 0.02 |
| Perfluorononanoic acid (PFNA) | ng/l | 0.18 | 0.081 | <LOD | 0.28 | 0.074 | <LOD | 0.11 | 0.031 | <LOD |  | 0.01 |
| Perfluorooctadecanoic acid (PFODA) | ng/l | 0.22 | 0.182 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.17 |
| Perfluorooctanesulfonamide (PFOSA) | ng/l | <LOD | <LOD | 0.07 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.05 |
| Perfluorooctanoic acid (PFOA) | ng/l | <LOD | 0.026 | <LOD | 0.045 | 0.018 | <LOD | <LOD | <LOD | <LOD |  | 0.01 |
| Perfluoropentanesulfonic acid (PFPeS) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 0.11 | <LOD | <LOD |  | 0.09 |
| Perfluoropentanoic acid (PFPeA) | ng/l | <LOD | 0.373 | 0.66 | <LOD | 0.301 | 0.65 | <LOD | 0.338 | 0.71 |  | 0.07 |
| Perfluoroundecanoic acid (PFUdA) | ng/l | <LOD | 0.094 | <LOD | 0.032 | 0.060 | <LOD | 0.033 | 0.045 | <LOD |  | 0.03 |

Notes:

- In April 2025, the sample at Otaci was not taken from the bridge but from a small inlet at the right bank a few hundred meters downstream.

|  |  | RAUT |  |  |  |  |  |  | ICHEL |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Bălţi |  |  | Ustia |  |  |  | Goian |  |  |  |  |
| Substance name | Unit | Apr | Jul | Oct | Apr | Jul | Oct |  | Apr | Jul | Oct |  | LOD ng/l |
| 1H,1H,2H,2H-Perfluorooctanesulfonic acid (6:2FTS) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 1.775 | <LOD | <LOD |  | 0.17 |
| 2,3,3,3-Tetrafluoro-2(heptafluoropropoxy)propionic acid (GenX) | ng/l | 0.065 | 0.24 | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.02 |
| Capstone A (CDPOS) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.04 |
| Dodecafluoro-3H-4,8-dioxanonate (ADONA) | ng/l | 1.703 | <LOD | <LOD | <LOD | 0.170 | <LOD |  | 0.157 | 0.164 | <LOD |  | 0.10 |
| Perfluorobutanesulfonic acid (PFBS) | ng/l | 0.74 | 0.593 | <LOD | 2.5 | 0.227 | 0.13 |  | 1.8 | 0.249 | <LOD |  | 0.08 |
| Perfluorobutanoic acid (PFBA) | ng/l | <LOD | 0.301 | <LOD | <LOD | 0.245 | <LOD |  | <LOD | 0.165 | <LOD |  | 0.11 |
| Perfluorodecanoic acid (PFDA) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | 0.109 | <LOD |  | 0.08 |
| Perfluorododecanoic acid (PFDoA) | ng/l | 0.031 | 0.035 | <LOD | 0.028 | 0.046 | <LOD |  | 0.051 | 0.042 | <LOD |  | 0.01 |
| Perfluoroheptanesulfonic acid (PFHpS) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.14 |
| Perfluoroheptanoic acid (PFHpA) | ng/l | 2.4 | <LOD | <LOD | <LOD | 0.133 | <LOD |  | 5.2 | <LOD | <LOD |  | 0.10 |
| Perfluorohexanesulfonic acid (PFHxS) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 8.3 | 0.084 | <LOD |  | 0.04 |
| Perfluorohexanoic acid (PFHxA) | ng/l | 0.07 | 0.081 | <LOD | 0.16 | 0.065 | 0.05 |  | 0.19 | 0.035 | <LOD |  | 0.03 |
| Perfluorononanesulfonic acid (PFNS) | ng/l | 0.070 | <LOD | <LOD | 0.026 | <LOD | <LOD |  | 0.13 | <LOD | <LOD |  | 0.02 |
| Perfluorononanoic acid (PFNA) | ng/l | 0.12 | <LOD | <LOD | 0.19 | 0.052 | <LOD |  | 0.60 | 0.015 | <LOD |  | 0.01 |
| Perfluorooctadecanoic acid (PFODA) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.17 |
| Perfluorooctanesulfonamide (PFOSA) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.05 |
| Perfluorooctanoic acid (PFOA) | ng/l | <LOD | <LOD | <LOD | 0.31 | 0.012 | <LOD |  | 0.015 | <LOD | <LOD |  | 0.01 |
| Perfluoropentanesulfonic acid (PFPeS) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.24 | <LOD | <LOD |  | 0.09 |
| Perfluoropentanoic acid (PFPeA) | ng/l | 3.6 | 0.190 | 0.75 | <LOD | 0.502 | 1.19 |  | <LOD | 0.335 | 0.41 |  | 0.07 |
| Perfluoroundecanoic acid (PFUdA) | ng/l | <LOD | 0.081 | <LOD | 0.042 | 0.072 | <LOD |  | 0.044 | 0.054 | <LOD |  | 0.03 |

|  |  | BIC |  |  |  |  |  |  | BOTNA |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Străşeni |  |  | Gura Bîcului |  |  |  | Chircăiești |  |  |  |  |
| Substance name | Unit | Apr | Jul | Oct | Apr | Jul | Oct |  | Apr | Jul | Oct |  | LOD ng/l |
| 1H,1H,2H,2H-Perfluorooctanesulfonic acid (6:2FTS) | ng/l | 0.904 | <LOD | <LOD | 2.615 | 0.89 | 1.00 |  | 0.422 | <LOD | <LOD |  | 0.17 |
| 2,3,3,3-Tetrafluoro-2(heptafluoropropoxy)propionic acid (GenX) | ng/l | <LOD | <LOD | <LOD | 0.22 | 0.22 | 0.03 |  | 0.45 | <LOD | <LOD |  | 0.02 |
| Capstone A (CDPOS) | ng/l | <LOD | <LOD | 0.06 | <LOD | 0.16 | 0.07 |  | <LOD | <LOD | <LOD |  | 0.04 |
| Dodecafluoro-3H-4,8-dioxanonate (ADONA) | ng/l | <LOD | <LOD | <LOD | <LOD | 0.261 | <LOD |  | <LOD | 0.255 | <LOD |  | 0.10 |
| Perfluorobutanesulfonic acid (PFBS) | ng/l | 5.0 | 0.298 | 0.36 | 1.2 | 0.212 | 0.13 |  | 4.6 | 0.287 | 0.32 |  | 0.08 |
| Perfluorobutanoic acid (PFBA) | ng/l | <LOD | 0.380 | <LOD | <LOD | 0.197 | <LOD |  | <LOD | 0.525 | <LOD |  | 0.11 |
| Perfluorodecanoic acid (PFDA) | ng/l | <LOD | 0.115 | <LOD | <LOD | 0.220 | <LOD |  | <LOD | <LOD | <LOD |  | 0.08 |
| Perfluorododecanoic acid (PFDoA) | ng/l | <LOD | 0.044 | <LOD | <LOD | 0.043 | <LOD |  | <LOD | 0.044 | <LOD |  | 0.01 |
| Perfluoroheptanesulfonic acid (PFHpS) | ng/l | <LOD | <LOD | <LOD | <LOD | 0.312 | <LOD |  | <LOD | <LOD | <LOD |  | 0.14 |
| Perfluoroheptanoic acid (PFHpA) | ng/l | 1.8 | 0.124 | 0.13 | 1.1 | 0.213 | 0.15 |  | <LOD | 0.209 | 0.17 |  | 0.10 |
| Perfluorohexanesulfonic acid (PFHxS) | ng/l | <LOD | <LOD | 0.08 | <LOD | 0.589 | 0.19 |  | <LOD | <LOD | 0.28 |  | 0.04 |
| Perfluorohexanoic acid (PFHxA) | ng/l | 0.31 | 0.226 | 0.17 | 0.56 | 0.185 | 0.08 |  | 0.42 | 0.113 | 0.05 |  | 0.03 |
| Perfluorononanesulfonic acid (PFNS) | ng/l | 0.028 | <LOD | <LOD | <LOD | <LOD | <LOD |  | 0.25 | <LOD | <LOD |  | 0.02 |
| Perfluorononanoic acid (PFNA) | ng/l | 0.45 | 0.082 | <LOD | 0.56 | 0.034 | <LOD |  | 0.28 | <LOD | 0.053 |  | 0.01 |
| Perfluorooctadecanoic acid (PFODA) | ng/l | 0.26 | 0.410 | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.17 |
| Perfluorooctanesulfonamide (PFOSA) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.05 |
| Perfluorooctanoic acid (PFOA) | ng/l | 0.42 | 0.065 | 0.10 | 0.63 | 0.019 | 0.06 |  | 0.017 | <LOD | 0.09 |  | 0.01 |
| Perfluoropentanesulfonic acid (PFPeS) | ng/l | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD |  | 0.09 |
| Perfluoropentanoic acid (PFPeA) | ng/l | <LOD | 0.822 | 2.04 | <LOD | 0.465 | 1.31 |  | <LOD | 1.026 | 1.79 |  | 0.07 |
| Perfluoroundecanoic acid (PFUdA) | ng/l | <LOD | 0.061 | <LOD | 0.077 | 0.082 | <LOD |  | <LOD | 0.076 | <LOD |  | 0.03 |

DIRECTIVE 2000/60/EC establishing a framework for Community action in the field of water policy

DIRECTIVE 2013/39/EU amending Directives 2000/60/EC and 2008/105/EC as regards priority substances in the field of water policy

For further details refer to: https://eur-lex.europa.eu/procedure/EN/2022_344

https://www.epa.gov/indoor-air-quality-iaq/what-are-volatile-organic-compounds-vocs

Related file can be downloaded via https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=celex%3A52022PC0540

Directive 2009/90/EC laying down, pursuant to Directive 2000/60/EC of the European Parliament and of the Council, technical specifications for chemical analysis and monitoring of water status

During another survey the Environmental Institute had to process 700 litres of water in order to achieve the required LOQ.

Screenshot from: https://www.meteo.md/images/uploads/Hydro/caracteristica/noiembrie_2025_ro.pdf

The LODs mentioned in this column were derived from the actual monitoring data.

Refer to the Final Report of the Environmental Institute for more obervations and recommendations regarding laboratory analysis.

Author’ additions: CAS number 50-29-3: 4,4′-DDT (p,p'-DDT); CAS number 789-02-6: 2,4′-DDT (o,p′-DDT); CAS number 72-55-9: 4,4'-DDE (p,p'-DDE); CAS number 72-54-8: 4,4′-DDD (p,p'-DDD)